## iRMA Glossary
Term | Definition | mark for filter | Source: | GlossaryExport-2020-05-14T16_00_18
.mc files | Files which are created and edited with the MEDIT tool.
.NET | Microsoft software framework for application development and execution.
2D | two-dimensional
3D | three-dimensional
3D Intervention | Adaptive 3D Intervention is a special examination mode for interventions under CT control like biopsies, drainages, RFA. The examining physician can trigger individual scans in the examination room at a valid table position. Images are displayed on a monitor inside the examination room. Needle entry point and needle target point can be defined. In that way, the examining physician can plan the movement of the surgical instrument precisely to avoid injuries of sensitive structures like vessels.
3D Reference Image | Is a project of a non-axial reconstruction series orthogonal to the direction of its image stack resulting in a single image.
3D-BP | 3D backprojector
3DVirtuoso | a Siemens workstation product
3DVV | 3D Vessel View
4D | 3D visualization of different points of time (e.g. beating heart)
4D | respiratory gated scans
4DMSPECT | a third party software (developed by Invia) for quantitive postprocessing of cardiac SPECT
A4CS | Adaptive 4D Cardiac Shuttle; Heart Perfusion Scan
A4DS | Adaptive 4D Spiral
AAL | Audio Algorithm Layer : This is used as a synonym for the ARMA library
AAMI | Association for the Advancement of Medical Instrumentation
AAPM | American Association of Physicists in Medicine
Absolute accuracy | Difference between measured value and its corresponding real value.
AC | Attenuation Correction
AC | Alternating Current
ACC | Accessories
Access control | User permissions will be checked before access to the data will be granted.
Access rights | Defined user permission - create, modify, read, delete, execute
Access to Radiology Information | The Access to Radiology Information Integration Profile specifies support of a number of query transactions providing access to radiology information, including images and related reports. Such access is useful both to the radiology department and to other
accessories | All equipment needed for positioning, fixation and stabilization of the patient. Includes also equipment neccessary for comfort and hygienic aspects, monitoring purposes (ECG leads, spiromters, etc)
accessories | all parts delivered and used with the machine in direct electrical and/or mechanical contact like matresses, head holders, foot switches, ECG monitors, spirometers et al.
ACE | Adaptive Communication Environment
ACK | Acknowledge
ACL | Audio Command Layer. This is a Audio SW layer, which implements the SW interface between the Gantry Control System and the AAL.
ACPI | Advanced Configuration and Power Interface convention Open standard that defines platform-independent interfaces for hardware discovery, configuration, power management and monitoring
ACQ | Acquisition system
ACQ-A | Acquisition system A
ACQ-B | Acquisition system B
Acquisition Modality | A system that acquires and creates medical images, e.g. a Computed Tomography scanner or Nuclear Medicine camera. A modality may also create Grayscale Softcopy Presentation States for the consistent viewing of images and create Key Image Notes.
Acquisition Phase Radiation | - Scanning: This phase begins as soon as the <Start key> is pressed. The phase contains the time span between measurement start and measurement end. A measurement starts with the start of its first range and ends with the end of its last range. The one or more ranges of a measurement are performed sequentially and each range has 3 execution phases (before radiation / during radiation / after radiation). - Before radiation: This is the first phase of a range and it comprises the start delay for the range. In this phase automatic table movements can be performed and the pre-range API (automatic patient instruction) is played. The phase ends with the start of radiation (first "x-ray on"). - During radiation: This phase covers the timespan from the "x-ray on" of the first scan of the range to the "x-ray off" of the last scan of the range. If the range contains multiple scans (e.g. sequence scan) the phase covers multiple "x-ray on" / "x-ray off" event pairs. - After radiation: This is the final phase of a range and begins as soon as radiation for the range is done. During this phase post-range API is played.
Acquisition Phases before radiation | - Measurement loading: This phase begins when the users presses the <Go Button> or when the system automatically loads the next measurement. During this phase the parameters for the next measurement are sent to the gantry and the IRS. The phase ends as soon as the measurement is accepted and successfully loaded at the gantry and IRS. - Preparation for start: This phase begins as soon as a measurement is loaded at the system and user actions are necessary before the measurement can be started. E.g. typically the user is asked to press the <Move Key> to position the table at the correct position (horizontally / vertically) or to tilt the gantry. The phase ends as soon as all start requirements are fulfilled. - Ready for start: This phase begins as soon as all start requirements for a measurement are fulfilled. In this phase the user is asked to press the <Start key> to begin the the measurement. The phase ends as soon as the <Start key> is pressed. This phase is automatically skipped for acquisitions which do not request a <Start key>, e.g. subsequent acquisitions in a Service workflow.
Acquisition Time | A pair of date and time stamps that defines the beginning of data accumulation for a particular CT image. Due to the multislice technology several CT images may have the same acquisition time (date and time). The corresponding DICOM attributes are Acquisition Date 0008,0022 and Acquisition Time 0008,0032.
ACR | American College of Radiologie
ACR-NEMA | American College of Radiology - National Electrical Manufacturers Association
ACS | Acquisition Control System (PET)
aCTiDAS | aCTivate Data Acquisition System
aCTiMAS | MAS as used in aCTivate platform
aCTivate | Program for development of M3 market CT platform
ACV | Adaptive Cardiac Volume
ACV | Adaptive Cardiac Volume, Mehrsegment-Rekonstruktion
ADC | Analog Digital Converter
Administration Portal | The Administration Portal is the central User Interface for the Service Engineer. It is based on the SY WebUI Framework. For Local Access, the Administration Portal is embedded in the clinical UI, together with other UI elements like Image segments or Position UI. For Remote access, the Administration Portal can be accessed through HTTP/HTTPS request and shows a reduced set of functions which is relevant for remote diagnostics without disturbing the customer´s workflow.
ADMIRE | ADvanced Modeled Iterative REconstruction
ADVANCED license | The term ADVANCED license indicates that an enhanced features within the individual task is either enabled or disabled (see also BASIC license).
AEC | Automatic Exposure Control
AERB | Atomic Energy Regulatory Board
AET | Application Entity Title (network node identification within DICOM)
AHA | American Heart Association
AHCI | Advanced Host Controller Interface
AI | Artificial Intelligence
AIAT | Assembling, Installation, Adjustment, Testing
AIF | Arterial Input Function
Air KERMA | KERMA in a given mass of air. The unit of air kerma is Gray (Gy).
AJS | CT Application Job Scheduler
Akü / Abb. | English / francais
AL | Nummer der deutschen Ausfuhrliste
ALARA | As low as reasonably achievable. Concept to optimize dose according to required image quality.
ALARP | As Low as Reasonably Possible
ALD | Arcing Location Detector
ALP | Arterial Liver Perfusion
Alpha Blending | Method for combining two images that uses both pixel colors and alpha values to determine the color of the resulting pixel. Alpha blending is used to display an image that has transparent or semi-transparent pixels. In addition to a red, green, and blue c
AMC | Angular mA Control
AMD | Short for Amendment, a formally published change document that modifies the content of an International Standard
AMPR | Adaptive Multiple Plane Reconstruction
AMU | Audio Master Unit : HW resources of the iMX6 CPU which is allocated for the AUDIO processing
Annual operation time | This is considered to be 3000 hours per year. This time is calculated as 365 days per year less 52 sundays less 12 holidays less 1 day for maintenance = 300 days * 10 hours per day = 3000 hours
ANSI | American National Standards Institute
Ao | Aorta
AP | Anterior-Posterior
AP | Angle Pulse - pulse signal triggered by an indicator at defined angular distances to divide a rotation into defined intervals for position interpolation
Apex | The Apex of the heart is the lowest superficial part of the heart
API | Automatic Patient Instruction
API | Application Programming Interface
aPK | aCTivate PreKit (basic fragment of the gantry)
APM | Asynchronous Programming Model
APP | Shortcut to mobile app - is a computer program designed to run on mobile devices
App | Application
Appointment | Time scheduled entry in the calendar view of the Organizer
appropriate | Ambiguous term. This term should be replaced with a list of the actions meant.
APR | Advanced Presentation State
AQP | Acquisition Preprocessing
AQWP | syngo Acquisition Workplace
ARC | Acquisition and Reconstruction Computer
ARI | Access to Radiology Information (IHE Profile)
ARMON | ARMA monitor. Command-line Tool on aCTimas for Controlling the AUDIO system ( like Command-line Tool on aCTimas for Controlling the AUDIO system ( like
ARS | Acquisition and Reconstruction System (PET)
AS | Ausführungsstand
AS | Abort Start
ASA | Advanced Smoothing Algorithm
ASAT | Audio SATellite board
ASCII | American Standard Code for Information Interchange
ASD | Acceleration Spectral Density
ASG | Anti-Scatter Grid
ASIC | Application Specific Integrated Circuit
ASPECTS | The Alberta stroke programme early CT score (ASPECTS) is a 10-point quantitative topographic CT scan score used in patients with middle cerebral artery (MCA) stroke. It has also been adapted for the posterior circulation.
ASSA | Application Specific Status Area
Asset Management | Asset Management provides a logical model of our customers infrastructure by identifying, controlling, maintaining and verifying the versions of Configuration Items (CIs) in existence e.g. Release contents, including component CIs and their version number
AT | Application Template
AT | Advanced Therapies
ATA / ATAPI | Advanced Technology Attachment with Packet Interface (interface standard for the connection of storage devices )
ATCP | Angular Tube Current Profile
ATO | Authority to Operate
ATTACHMENT | large binary object stored in the system and connected to an ELEMENT
ATTRIBUTETYPE | informationfield belonging to an ELEMENTTYPE
Audit Trail | Logfile to which each access to patient health information will be stored
Authentication | A user identifies him/herself at the system
Auto Range | Represents a special type of Measurement where two or more single Ranges that grouped and are executed directly one after the ather without pressign the START key in between.
Auto Tasking | Association of a via taskflow to a reconstruction fob (e.g. MM-Reading)
AVA | Advanced Vessel Analysis
AVAILABILITY | the period of time the RM-tool is available for users
AVATAR | An icon or figure representing a particular person in the virtual world
AVC | Adaptive Cardio Volume, in this document also referenced as "Bi Phases"
Average image | A postprocessed image of the same size as the input images where the value at each pixel (xi, yi) in HU is the average over the pixel values at that location in all input images. "Average image(xi, yj) = (Baseline(xi, yj) + Dynamic1(xi, yj) + Dynamic2(xi, yj) + ... Dynamicn(xi, yj)) / (n + 1)", for all i, j.
AVG | Average
AVI | Audio Video Interleaved
AVI | Audio Video Interleave
AWM | All Wire Material
AX | Angiography, Fluoroscopy and Radiography Systems
AX | Angiography, Fluoroscopy and Radiography Systems
BAC | Balancing Controller
Backend | Part of software responsible for processing information
Background Activities /Algorithms | The background activities are summarized / collapsed by a background element in the taskflow control. The background activities are executed automatically without user interaction. Background activities can be configured for a task in the protocol designer, too. The status of the background activities are displayed in the message blinds of the taskflow control.
Backprojection | Processing step of the CT reconstruction algorithm that processes parallel projections to images.
Backprojector | Hardware which does the Backprojection of the projections to an image.
Backup / Restore | Backup / Restore (BaRe) is a mechanism for saving the current configuration of a specific ICS in order to restore it after this ICS has been updated to newer Somaris/10 software version. This way, the system can be updated to newer software versions while retaining the current configuration on the system. (Term used by Syngo is Export / Import configuration).
BAM | Balancing Master
BAM2 | Balancing Master, 2-generation
bariatric | regarding obese patients
baseline | A baseline is a label for a set of requirements at specific versions.
BASELINE | a frozen set of ELEMENT versions in the system
Baseline image | Image used as baseline for pixel value and time calculations. Is a single (loaded) image or an image created by averaging several loaded images even from several series. The baseline image shows the state before a change in contrast enhancement starts. The acquistion time of the baseline image is the acquisition time of the first image used for the baseline image generation.
BASIC license | The term BASIC licence indicates that either the task is available or not (see also ADVANCED license).
Basic Security | IHE Integration Profile concerned with information security (includes HIPAA and similar aspects)
BAT | Band Assessment Tool
BB | Brush Block
BC | Business Component
BE | Back End
BER | Bit Error Rate
BEV | Beam’s Eye View
BHC | Beam Hardening Correction
BI | Bolus Injector
BI | Business Intelligence
BIC | Bolus Injector Control
Biograph PET | CT scanner of MI
BIOS | Basic Input Output System
BIS | Bureau of Indian Standards
BL | Business Line
BLD | Beam Limiting Device
Block | The received data is organized in packets. In this document such a packet is called a block. There are two different types of blocks: Header Blocks and Slice Blocks. Each received reading consists of a single Header Block followed by one or more Slice Blocks. The Header Block plus all Slice Blocks before the next Header Block are called a reading data structure.
BMD | Bone Mineral Density
BMI | Body Mass Index
BMP | Bitmap
BNC | BNC connector Short for British Naval Connector or Bayonet Nut Connector or Bayonet Neill Concelman, a type of connector used with coaxial cables such as the RG-58 A/U cable used with the 10Base-2 Ethernet system. The basic BNC connector is a male ty
Bolus | Injected media used to raise the contrast of blood vessels to other tissue.
Bolus Tracking | Process of real-time enhancement measurement to archive an optimal skip to the Diagnostic Range
BOM | Bill of Material
BP | Backprojector
BPM | Beats per Minute
BR | Brake Resistor
branch | Branching is the duplication of an object in such a way that the newly created object has initially the same contents as the version branched off from, and development can happen in parallel along both branches.
BRANCH | is the duplication of an object under revision control so that modifications can happen in parallel along both branches
BSA | Body Surface Area
BSM | Basic Scan Modes
BSS | Basic Service Set - Wireless LAN wireless network with access to the infrastructure via a single access point
BSSID | Basic Service Set Identification - Indication of the MAC address of the access point
Btn | Button
BU | Business Unit
BU CR | Business Unit Computed Tomography & Radiation Oncology
Bug-Fix | = SW Update
Bug-Fix | deleted entry! please use other entry!
BV | Blood Volume Mode
Bwd | Backward
BWG | Bildwiedergabegerät
BYOD | Bring your own device
CA | Conjoint Analysis
Ca | Calcium
CA | Computer Associates
CABG | coronary artery bypass graft
CAC | Coronary Artery Calcium
CAD | Computer Aided Diagnosis
CAD | Computer Aided Design
CaHA | Calcium Hydroxyapatite
CaliberRM | (purchased) Requirements Management Tool CaliberRM
Cali-MERO | Project Name (not used anymore); Project covers the customization of the purchased tool CaliberRM for H IM CT.
CAN | Controller Area Network
CANopen | CAN open bus, standard bus system based on CAN
CAP | Common Application Package
CAP Page | CAP component which contains the image segments
CAPA | Corrective Action, Preventive Action
CapGM | front end process of syngo based SW systems
CAPP | Customer Approval
CARE | Combined Applications to Reduce Exposure
CaSc | Calcium Scoring
Case | Patient examination, typically consisting of a series of images
CAT | Client Architecture Team
CAT | Computer Aided Tomography
CAT | Customer Approval Test
CATPHAN | Brand name of the currently available phantom for measurement of Low contrast detectability.
CAUD | Caudal
CAUSE U | the reason of a specific HAZARD
CB | Certified Body
CB | Certified Body
CBAZ | cone beam artefact z-direction
CBC | California Building Code
CBC | Cone Beam Correction
CBCT | Cone beam Computed Tomography
CBF | Cerebral Blood Flow
CBG | Generator Control Board
CBM | Continuous Bed Motion
CBR | Cone Beam Reconstructor (3D-Backprojector)
CBR D11 | Cone Beam Reconstructor, 3D-Backprojector Board, partno. 10187238
CBR-8G D12 | Cone Beam Reconstructor with 8GB Memory, 3D-Backprojector Board, partno. 10187239
CB-Report | Electrical Safety Report
CBT | Computer Based Training
CBTL | CB Testing Laboratory (IECEE)
CBTP | Controlbox Transfer Protocol. Proprietary Communication Protocol which is based on the SPDIF standard and which specifies the messaging between the COD and the CT MAS.
CBV | Cerebral Blood Volume
CC | Customer Complaint
CCC | China Compulsory Certification
CCC | Customer Care Center
CCEB | Coupled Coax Electronic Board
CCOAX | Coupled Coax
CCR | Channel Correction (data preprocessing step)
CCS | Cell Conditioning Sublayer, mixes (transmit direction) and splits (receive direction) cells into data and signal, performs the CRC check and generation. The CCS contains the PLI-Interface to the physical layer.
CCS | Cooling Control System
CCSC | Central Context Sharing Controller - a central ICWE component
CCSCR | Cell Conditioning Sublayer Control Register, part of the sublayers management interface
CCSSR | Cell Conditioning Sublayer Status Register, part of the sublayers management interface
CD | Compact Disc
CDI | Cell Data Interface, transports the data field of each cell and the connection information to the upper sublayer.
CDMS | Core Defect Management System Tooling for Defect Management based on TFS
CD-R | Compact Disc - Recordable
CD-ROM | Compact Disc Read Only Memory
CD-RW | Compact Disc - Rewriteable
CDS | Contouring Data Set
CDT | Contactless Data Transmission
CDT | Contactless Data Transmission, slipring for gigabit data transmission, based on high frequency fields.
CDT-IF | Contactless Data Transmission Interface on NPU
CE | Communauté Européenne
CE | Comunidad Europea
CE | Comunidade Europeia
CE | Conducted Emission
CEDARS | Cedars-Sinai is California's largest private hospital (Los Angeles, California) See also QPS and QGS.
Cedars Cardiac Suite | a third party software (developed by Cedars Sinai Medical Center) for quantitive postprocessing of cardiac SPECT
Cell | Atomic data unit, which is transported through the SiDaNet. All different data types are encapsulated in cells before they are transported. The switch uses the connection information, which is part of each cell, to guide the cell through the SiDaNet.
CES | Central Services
CFDA | China Food and Drug Administration
CFK | carbon fiber plastic
CFR | Code of Federal Regulations (USA)
CGI | Common Gateway Interface
Charge Posting | IHE Integration Profile with the objectives: - Standardize the Charge Posted Transaction to the Charge Processor - Reduces System Interface Installation Time - Between Clinical Systems and the Charge Processors - Reduces the Need of the Billing System to Have Knowledge of the Radiology Internals.
CHARM | Change Management Request
Chest Reading | see Pulmo 3D
CHG | IHE INtegration Profile Charge Posting
CHI | Common Human Interface
CI | CAN-Interface
CI | Cardiac Output Index
CI Board | Control Interface Board
CIC | Camera Integrated Computer
CIE | Commission Internationle de l'Eclairage / Commision on Illumination
Cine | Movie mode
CISPER | Limit class definition for EMC
CISPR | Comite International Special des Perturbations Radioelectriques (International Special Committee on Radio Interference)
Clinical Admin | A clinical user (e.g. a radiologist) with enhanced privileges (e.g. edit/save an Examination Protocol). A clinical administrator is also a clinical user and has additionally access to all features of a clinical user.
Clinical Task | A clinical task represents a clinically meaningful and self-contained piece of work performed by one user. Usually multiple clinical tasks are performed in a specific order to answer a clinical question. Clinical tasks may be assigned to different roles.
Clinical User | The clinical user is the standard user. Typical representatives of a clinical user are MTAs or doctors.
cloud computing | data processing is not done on local hardware but offered as a service via internet connection on remote hardware (server clusters).
Cluster | A Cluster is a set of Sequence scans of a Sequence Range which can be done within the patient`s breath hold time. Between two Clusters a pause is inserted to allow the patient to take a new breath.
CM | Contrast Medium
CM | Configuration Management
CME | Continuing Medical Education
CNR | contrast to noise ration
CNSS | Committee on National Security Systems
Co | Cardiac Output
COA | Certificate of Authenticity
COC | Collimator Control
COD | Control Device
CODIAG | Diagnostic Tool for computers
COEM | Customized Original Equipment Manufacturer
cog | center of gravity
cold start | system (gantry and console) is off, only DMS may be powered. Time counts, until system is ready for measurement with specified image quality.
combined system | a combination of a CT with another modality e.g. PET-CT or SPECT-CT
Comp On | Computers on (system power state with gantry power off)
Compact Series | Logical grouping of DICOM series objects
Component | A component inside the SiDaNet, which generates, receives and distributes data traffic.
Con_Box | Control Box
CON_GPC | Controller Task on aCTimas, which contains the panel and key control on the CT Gantry
CON_SAF | Controller Task on aCTimas, which control safety functionality at CT system
CON_SPC | Controller Task on aCTimas ( previous name : CON_MAS), which controls the AUDIO system
ConD | Control Device
configure | change program parameters via a resource file or similar (maybe service only)
configured test case | A test case in the context of a configuration.
Confirmation | Last step in a bi-directional communication interaction where the feedback from the destination device is received by the source device.
Connection | A way through the SiDaNet between a data source and a data sink. Data could only be transported after establishing a connection between the two involved devices. Multiple connections could share one link.
Consistent Presentation of Images | The CPI Integration Profile specifies a number of transactions that maintain the consistency of presentation for grayscale images and their presentation state information (including user an notations, shutters, flip/rotate, display area, and zoom). It also defines a standard contrast curve, the Grayscale Standard Display Function, against which different types of display and hardcopy output devices can be calibrated. Thus it supports hardcopy, softcopy and mixed environments. The systems included in this profile are hospital-wide and radiology-department image rendering systems such as: - Review or diagnostic image softcopy display stations (stand-alone or integrated with a HIS, RIS or PACS) - Image management and archiving systems (PACS) - Hardcopy image producing systems on various media such as film or paper - Acquisition modalities
Console | Image Control System
Content Work Items | All Work Items stored in the database that hold project relevant data. (e.g. Requirements, Test Cases)
Contrast Delay phase | One step in the injection protocol when fluid is not dispensed.
Contrast protocol | Part of an Examination Protocol, describes the contrast relevant parameters, e.g. CM type
Contrast scan | A scan protocol with a Contrast protocal (see also Native scan)
Control Box | Control box is part of the CT system with HW keya to control the system, e.g. MOVE
ControlArea | The control area guides the user through the examination. It consists of two regions. The upper part is called TaskflowControl, the lower one ExecutionControl.
COPD | Chronic Obstructive Pulmonary Disease
CorRea | Correct and Rearrange
Corresponding Topo | A list of corresponding Topogrmas for a Tomogram is found: 1. Skip all Tomograms right before (in the order of the protocol) of the Tomogram. 2. From there take all Topograms before the next tomo which are in the same FoR. Any Topogram from this list is a corresponding Topo.
COTS | Commercial off-the-shelf
Coupled mode | The Coupled mode is activated, if a scan range is synchronized with the connected bolus injector.
CP | Control Protection
CP | Content Page
C-path | Control Path
CPI | Cardio Pulmo Interface
CPIN | Cos Phi Inductor
CPP | Control Protection Protection
CPR | Computerized Patient Record
CPR | Curved Planar Reformat
CPT | Contactless Power Transmission
CPU | Central Processing Unit
CR | Change Request
CR | Computed Tomography & Radiation Oncology
Cradle | combination of jacket and hook
C-RADS | The CT Colonography Reporting and Data System (CRADS) project was a collaborative effort that developed terms and guidelines for CT Colonography with categories for describing colonic findings.
CRAN | Crandial
CRC | Cyclic Redundancy Check
CRC | Cyclic Redundancy Check, checksum over a number of bytes to check the integrity of the bytes after reception.
CRM | Customer Relationship Management
Cron | Cron job is the time-based job scheduler in Unix-like computer operating systems e.g. Microsoft server technology. The name cron comes from the Greek word [chronos] for time
cRSP | Common Remote Service Platform
CS | Cancel Start
CS | Customer Services
CS_MHC | Customer Services Message Handling Concept
CSA | Canadian Standards Association
CSA | Common Software Architecture
CSE | Customer Service Engineer
CSI | Cell Sync Interface, Interface to synchronize the outgoing data traffic to the data rate of the incoming traffic. This interface is independent from the concrete implementation of the physical layer.
CSM | Central Strategy Management
CSM | Compatibility Support Module
csv | A comma-separated value text file (sometimes called comma-delimited files), commonly used to exchange files between database systems that use different formats.
CT | Computed Tomography
CT image | Original or calculated image from CT modality, that contains pixel information.
CT values | special unit of pixel values in CT images where 0=water and -1000=air
CTA | CT Angiography
CTAC | CT image used for MI’s Attenuation Correction
CTDI | Computed Tomography Dose Index
CTDI | Volume Computed Tomography Dose Index, in mGy according to IEC 60601-2-44. An estimate of the average dose over the scanned volume in a CT dosimetry phantom that is used to allow comparison of performance against a reference dose value set for the purpose of promoting optimization of patient protection.
CTDI VoI | CT Dose Index Volume
CTDI100 | Computed Tomography Dose Index 100, CTDI value from -50mm to +50mm
CTDIfree air | Computed Tomography Dose Index, CTDI value in the absence of any dosimetry phantom and patient support
CTDIvol | CT Dose Index Volume, CTDI value weighted and pitch factor corrected
CTI | Comparative tracking index
CTV | Clinical Target Volume
CTWP | CT Workplace
CU | Central Unit to supply, manage and diagnose the Rx/Tx-modules
CUDA | Unified design architecture
Curve Image | An image which does not contain any visible pixel information but shows tabular or graphic results. If the image does not have any tabular and graphic results it is termed as "empty curve image".
Custom Pack | Custom packs are software distribution kits that may be used for short-term customer specific solutions.
CUT | Clinical Use Test
CV | Business Unit Components & Vacuum
CVI | CARE Vision
CX | Circumflex
D2I | Dedicated2Imaging LLC, Portsmouth, USA
D32 | CIB CAN Interface Board
DaP | Data access Protocol
DAS | Data Access Sublayer, delivers the interfaces to access the data types, based on the cells data via software or hardware. The different supported data types are variable, memory and packet. Separates the data to the different hardware and application (software) interfaces.
DAS | Data Acquisition System
DASCON | Data Acquisition System Controller
DASCR | Data Access Sublayer Control Register, part of the sublayer’s management interface.
DAS-CRC | CRC calculated for the contents of a block including the footer for transmission from DAS to Receiver
DASSR | Signal Access Sublayer Status Register, part of the sublayer’s management interface.
DASTMRCR | Data Access Sublayer Transmit Memory Receive Control Register, part of the sublayer’s management interface.
DAT | Digital Audio Tape
Data Consistency | Ensure that the relation between all single data items in a data set is correct
Data Integrity | Ensure that all single data items in a data set belongs together
Data Readings | Readings of this type contain valid measurement data or test data.
DataRole | DataRole syngo.via specific data model clustering DICOM data
dB | Decibel
dBA | A-weighted decibels
dBFS | Decibels relative to full scale. Measures decibel amplitude levels in digital systems which have a defined maximum available peak level.
DBG | DrosselBauGruppe
DC | Direct Current
DCI | Data Communication Identifier. Specifies the connection of a cell.
DCI | Data Communication Identifier. Specifies the connection of a cell. Device SiDaNet component, which generates and consumes data traffic. Each switch contains exactly one network port.
DCON | System Controller with SiDaNet Device (SPECT)
DCR | Drive Control Rotation
DDK | Driver Development Kit
DDT | Data Distribution Tag, information on which ports and in which timeslot(s) a cell must be sent.
DE | Dual Energy
default value | A pre-defined value for a specified attribute
DELETE | erase of all information from the system
Delivery Delay | The difference between Time Until Images and Recon Duration (may also include the delay starting a recon job)
DeltaXML | DeltaXML is a XML-Diff tool which finds changes between two versions of a XML-file.
DESCRIPTIONFIELD | text field with special properties concerning compfortable formatting.
DESS | Dual Energy Single Source
Detail Planning | Generation and Review of test plan, i.e. definition of all test cases to be executed and not to be executed. The different system types are taken into account.
Detector Channel | Data output signal from the DMS.
Detector row | Physical row of the detector, produced by subdivision of the detector
Device | SiDaNet component, which generates and consumes data traffic
DFS | Dynamic Frequency Selection
DHHS | Department of Health and Human Services
DHR | Device History Record
DHTML | Dynamic Hypertext Markup Language
DI | Dependency Injection
Diagnostic Scan | The scans used for acquiring images for diagnostic purpose. Topograms, Premonitoring, Monitoring and Testbolus scans are not considered to be diagnostic scans.
Diagnostic time | Diagnostic time is the time from positioning the patient in the systems isocenter until the diagnostic information is shown on the diagnostic tool.
DIC | Data Image Chain
DICE | Dice Coefficient (metric for validation of medical volume segmentations)
DICOM | Digital Imaging and Communication in Medicine
DIM | Data Input Multiplexer, splits the cell data of CDI-Interface to the receive functions for each data type.
DIN | Deutsches Institut für Normung (German Institute for Standardization)
DiP | Display Protocol (configuration file which specifies the way images are displayed)
DIP | Dual in-line package
DIP-switch | package of switches on/off, assembled to electronic boards to perform basic configurations
DIREC | Dual Input Receiver (2 parallel LWL input channels)
DiRec D26 | Dual Input Receiver for PCI-X, part number 8098118
Direct Postprocessing | Immediate postprocessing after the last scan of examination, defined in the examination protocol
Directory Service | Mechanism to share central stored information network wide.
DIRex D27 | Dual Input Receiver for PCI-Express, part number 8098613
DIRexSim | Simulator Board; counterpart for DIRex D27 in the test environment
DIS | Digital Imaging System
DISF | DISplay Front
Display States | All Parameters and UI controls could have the following states: enabled: the parameter / UI control is visible and editable disabled: the parameter / UI control is visible but dimmed invisible: the parameter / UI control is not visible on the UI read-only; the parameter / UI control is visible but always disabled
DISR | DISplay Rear
DLL | Data Link Layer, communication layer described inside this document.
DLL | Dynamic Link Library
DLLMI | Data Link Layer Management Interface, defines the access for control and status registers.
DLP | Dose Length Product
DMA | Direct Memory Access
DMAR | DAS (Data Acquisition System) - Controller with MAR (Master Rotating) function
DMC | Data Measurement Control
DMI | Desktop Management Interface
DMR | Device Master Record
DMS | Data Measurement System
DMWL | Dicom Modality Worklist Query
DOA | Defective on Arrival
DOB | Date Of Birth
Dock | station for charging and storage of mobile device(s) with external power supply
Document | One or multiple Work Items and / or Test Results which can be seen as a unit, and are required to be saved as a Document. (e.g. TestLog, Review Unit) (Can be used for Reviews, Releases and Test Result LeadOuts to SAP)
DoD | Department of Defense
DoDI | Department of Defense INSTRUCTION
DOM | Data Output Multiplexer, merges the transmit functions in the transmit direction of the DAS to the CDI-Interface.
DOM | Dose Modulation
Dose modulation | Dose modulation with the target to reduce the dose and keep the image quality.
downtime | the time the system cannot be used for any examination due to a defect within the system
DP | DisplayPort
DPMS | Display Power Management Signaling
Drag and Drop | Drag-and-drop is the action of (or support for the action of) selecting a virtual object by "grabbing" it and dragging it to a different location or onto another virtual object.
DRR | Digitally Reconstructed Radiograph
DRS | Dedicated Reconstruction System (SPECT)
DS | Dual Source
DS | Design Specification
DSA | Digital Subtraction Angiography
DSCT | Dual Source Computed Tomography
DSDE | Double source dual energy
DSF | Data Switching Function of a switch, analyzes a received cell and generates the information, at which ports and when the cell must be transmitted.
DSF | Data Streaming Frontend
DSM | Data Streaming Master
DSS | Device Synchronization Sublayer, is responsible for the synchronization of the transmit data rate to the receive data rate inside a device. It always works as a slave.
DSS | Data Streaming System
DTD | Document Type Definition
DTS | downtime system = downtime hours for partial and total failure / operating hours in %
Dual Topo | Combination of two Topograms with different tube positions (A.P. and LAT).
Dummy Reading | Readings of this type are for error checking purposes only and do not get stored.
DuSo | Dual Source
DVD | Digital Versatile Disc
DVH | Dose Volume Histogram
DVI | Digital Visual Interface
DVI-A | Digital Visual Interface - Analog
DVI-D | Digital Visual Interface - Digital
DVI-I | Digital Visual Interface - Integrated
DVT | Design Verification Test
Dw | Water Equivalent Diameter
DWIT AS | Defect Work Item Type – Aligned Solution
DXS | Data Switching Sublayer, switches the received data of each cell to the transmit ports. The criteria for the selection of the transmit ports is the connection information contained in each cell.
DXSCR | Data Switching Sublayer Control Register, part of the sublayers management interface.
DXSDCR | Data Switching Sublayer Control Register, part of the sublayers management interface.
DXSPSR | Data Switching Sublayer Port Selection Register, part of the sublayers management interface.
DXSSR | Data Switching Sublayer Status Register, part of the sublayers management interface.
Dynamic image | In the context of this document an input image of a particular slice position of a patient body taken sometime after a change in contrast enhancement started
Dynamic loading | An application can be started from the PB or from the SynGo 2D Viewer by selecting the Application under the "Patient" (in later version "Postprocessing") menu entry. It can also be started by pressing the corresponding application button in the control area of 2D Viewer or from toolbar button in PB. If images were selected in PB or 2D Viewer before the application is started, these images are automatically loaded into the application at start-up.
Dynamic Multiscan | Scanning technique of continuous data acquisition of multiple scans over single table position. (see "Multiscan")
dynamic scan cycle time | time from one 360° scan to the next one with fastest rotation speed, asynchronous reconstruction and no table feed.
Dynamic Test Configuration | This type of Test Configuration is defined by a list of Standard Test Configurations. Dynamic Test Configurations can be either EACH or ANY. During Test Planning a Dynamic Test Configuration is translated to a single or multiple Configured Test Cases, depending on the EACH and ANY.
e.g. | exempli gratia (for example)
e.soft | Software platform for display and processing of SPECT images and for controlling of SPECT and CT examination
e.via | embedded syngo.via
E_Box | Electronic Box
EAB | System Flow Chart
EBC | Emergency Brake Control
EBT | Electron Beam Tomography
ECC | Error Correction Code
ECC | Elliptic Current Control
ECC | Elliptical current control
ECG | Electrocardiogram
ECG | Electrocardiogram; The ECG will not be used for diagnostic purposes or monitoring.
ECG-gating | Retrospective ECG-based selection of raw data for reconstruction
ECG-pulsing | Cardio specific dose modulation
ECG-triggering | Prospective ECG-controlled release of X-ray in sequence scanning
ECS | Embedded Control System (CT Product name)
ED | Evidence document
ED | Electron Density
ED | end diastolic volume
EDA | ECG controlled Dose Automation
EDO | Eigth Detector Offset
EE | Electrical Engineering
EEC | European Economic Community
EED | End to End Delimiter, used for the coordination of data exchange between the source and the destination.
EEPROM | Electrically Eraseable Programmable Read Only Memory
EF | extra-focal
EFC | Extra Focal Radiation Collimator
eFoV | extended Field of View
EFR | extra focal radiation
EFS | extra focal radiation
EFTA | The European Free Trade Association
EHD | Ethernet Highway Driveup
EHR | Engineering History Record
EHS | abbreviation for environmental protection, health and safety
EHSD | Siemens EHS Document for Product-Related EHS Protection
EI | Evaluation Index
EID | Energy Integrating Detector
EIRP | Equivalent Isotropic Radiated Power
Electron Density | Electron Density needed in radiation therapy to calculate the beam absorption of the high energy beam based on CT absorption values (= HU) based on low energy images.
ELEMENT | a data record in the RM-tool (e.g. a requirement)
ELEMENTTYPE | Abstract definition of all information belonging to a kind of an object.
ELH | Requirement Specification
ELV | Extra Low Voltage
E-Manual | Consist of e-manual executables, e-manual config data and e-manual content
E-Manual configuration | Configuration data for e-manual executables
E-Manual content | Customer documentation data provided for e-manual
E-Manual executables | Viewer & setup applications with related skripts for the e-manual
EMAS | Eco-Management and Audit Scheme
EMC | Electromagnetic Compatibility
EMCY | CANopen emergency message
emergency start | cold start without checkup. Limited image quality is accepted.
EMI | Electromagnetic Interference
eMMC | embedded multi-media card
EMR | Electronical Medical Record
EMW | Embedded Middleware
EN | European Norm, European Standard
encore | Unit of combined SPECT and CT, product name
ENDO | left ventricular endocardial border
EP | Enterprise Platform
ePDF | Encapsulated PDF
EPH | Design Specification
EPI | left ventricular epicardial border
EPO | Emergency Power Off
EPR | Electronic Patient Record
EPRC | Electronic Product Radiation Control
EPROM | Erasable Programable Read-Only Memory
EPS | enhanced productivity service
ERS | Engineering Requirement Specification
ES | End Systole
ES | Erzeugnisstand
ESC | Escape key
ESC | European Society for Cardiology
ESD | Electrostatic Discharge
ESE | Ethernet Switch External
ESI | Ethernet Switch Internal
ESS | Extended Service Set - Wireless network based on several overlapping basic service sets (BSS)
ESSID | Extended Service Set Identifier
eSW | embedded Software
EUV | Vertrag über die Europäische Union
EVA | Evaluation Function
evaluation | a complete evaluation of a patient, if status handling of the application is possible, an evaluation can continued in another session
EVD | IHE Integration Profile Evidence Documents
Eventlog | Repository to store the messages (set by the components)
Evolve | Evolve is a Siemens Medical program which includes hard- and software updates and upgrades.
EVOLVE Syngo | Replacement of the computer system to ensure state-of-the-art hardware and software after a defined period of time for customers with EVOLVE contracts. It includes syngo, too.
EWG | Europäische Wirtschaftsgemeinschaft
Examination | Scan Process and Reconstruction process and patient handling process (incl. service)
examination procedure | examination procedure comprises all steps and aspects of an exam from patient data registration, scanning, reconstruction to post-processing, diagnosis including reporting and distributing the information.
Examination Protocol | Scan Protocol with (optional) links to contrast protocols and clinical decision trees.
Exposure time | Time from x-ray on until x-ray off
extended message text | the extended message text provides the user with detailed information for each message
External Network | Ethernet connection between the scanner and external components (IES, camera, etc.)
Extranet | Special form of Internet with limited access and features to prevent the system against illegal access and virus (fire wall).
F H C M | Item of risk analysis: Function / Hazard / Cause / Measure
FAILOVER | redundancy within a computer network in order to ensure network availability after the failure of a particular piece of hardware
FAST | Fully Automatic Scanner Technology
FastLink | Bidirectional multiplexed serial link operating at 1.25 GBit/s
Fastlink | Proprietary communication protocol on glass fiber
FAT | File allocation table (MSDOS file system)
FB | Function Block
FCC | Federal Communications Commission
FCS | Ethernet Frame Check Sequence
FCT | Frequency Converter Tilt
FD | Findings Dialog
FDA | Food and Drug Administration (USA)
FE | Front End
FFS | Flying Focal Spot
FFT | Fast Fourier Transformation
FGUM-KEY | a unique trace key assigned to a specific measure
FIFO | First In First Out
FIL | Filament Control
file-link | A link to a file, located at the filesystem. By clicking this link, the specified file opens.
Fine Plan | The Fine Plan contains all configured test cases which are assigned to a baseline. In the Fine Plan the configured test cases are assigned to Test Runs where they shall be tested.
FIPS | Federal Information Processing Standard
Fix-axial and non-fix-axial reconstruction / recon | A fix-axial reconstruction is a reconstruction where the flag FixAxial is turned on. In this case the FoV is quadratic, the ReconAxis is aligned to the z-axis, no 3D Reference Image is created. A non-fix-axial reconstruction is a reconstruction where the flag FixAxial is turned off.
Fixed Image | Reference images during registration.
Fixed Image Shadows | All images sharing same FoR with the Fixed image.
FKM | Forschungskuratorium Maschinenbau (FKM) e.V.
Floating Image | Model images during registration.
Floating Image Shadows | All images sharing same FoR with the Floating image.
FLU | Fluoroscopy
Flyout | Extended tooltip which is sensitive to the mouse (e.g. can be clicked). The dynamic behaviour is the same as for ordinary tooltips: Opened when hovered over by the mouse, closed by leaving the element.
FMEA | Failure Mode and Effects Analysis
FN | Findings Navigator
FOB | Free on Board
FOC | Focus Position Control
FOL | Fiber-Optic Link
Footer | A structure at the end of each block containing transmission protocol information: "Reading Number", "Slice Number" plus status information.
FoR | Frame of Reference
FoV | Field of View
FPA | Floating Point Arithmetic
FPE | First Pass Enhancement
FPGA | Field Programmable Gate Array
FPGA-B | FPGA on CBR used for Backprojection
FPGA-C | FPGA on CBR used for Control
FPI | Frame Position Indicator, defines the position of a cell inside a frame. The FPI is used for signal multiplexing and QoS management of data.
FPID | Function Point ID, ident for action by application
FPOD | Fast Planning Organ Detection
FQ | Full Quality
Frame of Reference | Definition of a set of images which belong together, e.g. same patient position (see also DICOM definitions).
Frame Set Planning Basse | The graphical planning of Recon Jobs on the Reconstruction Task is supported by using a planning base. There are two different types of planning base: Frame Set Planning Base and Volume Planning Base. A Volume Planning Base must fulfill two conditions: All frames are on a different position and all image positions are in a straigth line. Otherwise it is a Frame Set Planning Base. The planning segments (axial, coronal and sagittal) show three orthogonal planes of the image volume for a Volume Planning Base. For a Frame Set Planning Base the axial planning segment shows the image stack. In CT environment all reconstruction series are of type Volume Planning Base except for multi-phase reconstructions which are covering the same volume multiple times. For technical details see http://sdn.healthcare.siemens.com/, search: ´DefaultVolumeValidator´.
f-Role | functional data role
FRU | Field Replaceable Unit
FRW | Frame Work
FS | Functional Specification
FSB | Frontside Bus
FSM | Foot switch (movement)
FSP | Functional Specification
FSX | Foot Switch for starting Xray
FT | Function Test
FTP | File Transfer Protocol
FuLi | Funnel Light
FUNCTION F | a specific solution / functionality in the Computed Tomography and Radiation Oncology environment
Function point ID | Parameter in an audit entry which identifies the application specific action being audited. (e.g. "ImageDelete")
FVFR | first visit fix rate
FW | Firmware
Fwd | Forward
FWHM | Full Width at Half Maximum
GaLi | Gantry Lighting
Gantry | main part of the CT Scanner containing the rotating drum.
gated | same for PET and SPECT as multi-phase for CT: multiple volumes scanned during the same examination but at different time points during the R-R interval
GB | Gigabyte (10^9 Bytes)
GB | Guobiao, chinesisch für "Nationaler Standard"
GB 18030 | Chinese government standard describing the required language and character support neccessary for software in China
GB/s | Gigabyte (10^9 Byte) per second
GB/T | Guobiao/tuījiàn, chinesisch für „empfohlener nationaler Standard“
GBit/s | Gigabit (10^9 Bit) per second
G-Box | Generator Box
GCF | Gantry Configurator
GCM | Gantry Control Message
GDL | GREC-FL Driver / Library
GE | Greater than or Equal to
GEN | Generator
GFW | Protection Module Generic FW
GGN | Ground glass nodule
GIA | Graphical Indicator Area
GiB | 2^30 Bytes (Gibibytes)
GiB | Gibibyte (2^30 Bytes)
GIM | Gantry Interface Module (PET)
GLOSSARY | composition of multiple term definitions and GLOSSARY ENTRY
GLOSSARY ENTRY | definition or abbreviation of a term that needs to be described
GM | Gantry Mechanics
GMP | Good Manufacturing Practices
GND | Ground
Gordon Unit | Scale unit of logarithmic detector signal
Gordon Unit | Logarithmic scale to define a ratio: x[GU] = 512 * ln(I_1/I_0) / ln 1.25
GOST R | Certificate of Conformity for russian market
GPA | Gantry Panel Assembly
GPB | Gantry Panel Buttons
GPC | Gantry Panel Control
GPI | Gantry Panel Interface
GPIO | General Purpose Input Output
GPT | Gantry Panel Touch
GPU | Graphic Processing Unit
GR | Generic Reader
GREC | GBit Receiver Board D24, predecessor of GREC-FL
GREC-FL D25 | Gbit receiver board with Fast Link Control Interface, partnumber 7742104
Growing Topogram | Topogram image where the image display is on-line with measurement.
GRP | Graphical Reconstruction Planing
GRP | Graphical Range Planning
GSP | Graphical Slice Positining
GSV | Gantry Server
GT | Ground Truth
GTC | Gantry Translate Controller
GTIN | Global Trade Item No.
G-Tree | Hierarchical order of DICOM data which sorts the data with regard to different aspects like modality, studies, examination time and identifies valid volumes.
GTV | Gross Tumor Volume
GUI | Graphical User Interface
GUID | Globally Unique IDentifier
Guidance System | Laser guidance for Interventional CT, e.g. Biopsy.
Gy | Gray Unit
hanging protocol | The Hanging Protocol composite IOD (information object definition) contains information about user viewing preferences, related to image display station (workstation) capabilities. The associated Service Classes support the storage (C-STORE), query (C-FIND) and retrieve (C-MOVE) of Hanging Protocol Objects between servers and workstations. The goal is for users to be able to conveniently define their preferred methods of presentation and interaction for different types of viewing circumstances once, and then to automatically lay out image sets according to the user’s preferences on workstations of similar capability.
HAV | High Attenuation Volume
HAZARD ¦ IMPERILMENT G | potential source of harm
HBC | Half beat cardiac
HBR | Heart beat rate
HCE | High contrast edge-enhancing filter
HD | Hard Disk
HD FoV | High Definition Field of View
HDC | Hybrid Data Coding
HDD | Hard Disk Drive
Header Start | A control in the received data stream indicating, that the next block received will contain header information of the next reading. There are three different types of header start control, one for dummy reading, one for data readings and one for resend data readings.
HF | High Frequency
Hibernate | System power state with reduced power consumption (standby)
HIC | High Integrity Component
HIPAA | Health Insurance Portability and Accountability Act
HiRes | High Resolution
HIS | Hospital Information System
HISTORY-RELEVANT DATA | stored information for every ELEMENT in the system when content was changed. (date, author/editor, type and size of modification...)
HL7 | Health Level 7, protocol in RIS environment
HLA | Horizontal long axis
HMI | Human-machine interface
HNG | Healthineers Needle Guidance
HNGI | Healthineers Needle Guidance Interface
HomeScreen | The HomeScreen is the start page of Somaris/10. By clicking on one of the show buttons the user van choose a workflow. The available options are: - Organizer - Service Workflow - Prepare a patient - Scan or Reconstruction Workflow
Hook | charger and holding device for jacket (with tablet-PC)
HP | High Power
HP2 | P82A (high power, 2 cm)
HP4 | P82B/C(High power,4cm)
HPI | Hepatic Perfusion Index
HPS | High pitch spiral
HPT | High Power Transmission
HQ | Headquarter
HR | Heart Rate
HRV | Heart rate variability
HSC | Headquarter Support Center
HTML | HyperText Markup Language
HTP | Horizontal Table Position
HTTP | Hypertext Transfer Protocol
HU | Hounsfield Unit
HV | High Voltage
HVL | half-value layer
HVT | High Voltage Transformer
HW | Hardware
Hybrid System | Combination of a CT system with another imaging modality, e.g. PET, SPECT or MR, for the purpose of combining functional and anatomical information
hz | Hazard
I2C | Inter-Integrated Circuit
I2S | Serial Audio transmission HW protocol/format
IB | Installed Base
IBC | Internet Based Connectivity
IBHC | Iterative Beam Hardening Correction
IC | Industry Canada
ICA | Ion Catcher
ICG | Inter Cell Gap, distance between two cells to add additional synchronization information into the data stream and to balance different transmit and receive baud rates.
ICS | Image Control System
ICS SW | Scan Application running on the ICS
ICT | Interventional Computed Tomography
iCUT | iCUT internal Clinical Use Test
ICWE | Integrated Cardiac Workplace Environment
ID | Identifier
IDL | Intelligent Display Layout
ID-PROM | Identifier Programmable Read Only Memory
IDT | Integrated Debug Tool
IDT | Integrated Diagnostic Tool: in circuit diagnosis tool of SiDaNet
IEC | International Electrotechnical Commission
IECEE | IEC System of Conformity Assessment Schemes for Electrotechnical Equipment and Components
IECS | Integrated Electonics Box Site connection
IEEE | Institute of Electrical and Electronics Engineers
IES | Image Evaluation System
IES2.0 | 2nd syngo.via client as separate workplace as part of IPS
IF_UMAS | InterFace Universal MAster Stationary
IFD | Interactive Finding Dialog
IFD | Interactive Findings Dialog
IfU | Instructions for Use
IGBT | insulated-gate bipolar transistor
IGRT | Image Guided Radiation Therapy
IHE | Integrating the Healthcare Enterprise The IHE initiative is a project designed to advance the state of data integration in healthcare. Sponsored by the Radiological Society of North America (RSNA) and the Healthcare Information and Management System
IHE | IHE: Integrating the Healthcare Enterprise The IHE initiative is a project designed to advance the state of data integration in healthcare. Sponsored by the Radiological Society of North America (RSNA) and the Healthcare Information and Management System
IIS | Internet Information Server
IIW | International Institute of Welding
IKM | Image and Knowledge Management (Healthcare Business Unit)
ILG | Integrated Laser Guidance
iLO | Integrated Lights-Out, or iLO, is an embedded server management technology exclusive to Hewlett-Packard, similar in functionality to the Lights out management (LOM) technology of other vendors. iLO is useful as a Ethernet technology for servicability issues.
Image Creator | A system that creates additional images and/or presentation states and/or Key Image Notes and transmits them to an Image Archive. It also makes requests for storage commitment to the Image Manager for the data previously transmitted and generates Performe
Image Display | A system that offers browsing of patients’ studies. In addition, it supports the retrieval and display of selected sets of images, presentation states and Key Image Notes.
Image mode | In image mode filming, the transfer of the images to the camera takes place image by image.
Image Segment Layout | It focuses on the arrangement, size and number of displayed images segments. It is embedded within the Task Area Layout.
iMAR | Iterative Metal Artefact Recuction
iMAS | MAS from the CT Baukasten, compatible to the aCTivate and impact platform
ImpaCT | Program for development of a M1/M2 market CT platform
IMS | Imaging System (IRS + ICS)
IMS | Imaging System (IRS+ICS+IPS)
IMS | Imaging System (IRS + ICS + IWP)
IMS | Imaging System (ICS + IES + IRS)
iMX6 | The iMX range is a family of Freescale Semiconductor proprietary microprocessors. The iMX6 is used on the aCTiMAS.
Indication | Second step in a communication interaction where the reception of data is indicated at the destination device.
Initial Planning | Test Center creates and adjusts the initial planning concerning personnel, systems, content and schedule for all tasks and roles.
Injection phase | One step in the injection protocol when fluid is dispensed
Injection Phase Element | Rectangle representing injection phase (contrast / saline) on the <Timeline>
Injection protocol | Sequence of configured <Injection phases> and <Delay Phase>s
Injector LocalProcedure | A workflow at injector to prepare the main injection procedure, e.g. KeepVeinOpen, TestInjection.
INMA | INtercom MAster
INMETRO | National Institute of Metrology, Quality and Technology (Brazil)
Input image | The baseline image and the dynamic images comprise the set of input images.
Instant 3D | Reconstruction of 3D images within examination procedure
Internal Network | Ethernet connection between the components of the scanner
Internet | Electronic information service, e.g. E-mail
Interrupt Events | User Interrupt: Stop STOP key on control box pressed Suspend SUSPEND key on control box pressed Cancel <Cancel> on UI clicked System Interrupt: Abort FW or HW error in IRS or gantry Error internal error Note: For successfully finished scans the state is "Success".
IO | Input / Output
IoC | Inversion of Control
IOD | Information Object Definition
IP | Internet Protocol Link A wireless, wire or fiber point-to-point connection between two SiDaNet ports. Each links establishes a bi-directional, full-duplex connection. A bandwidth of a bi-directional link could be increased by a additional unidirectional link in the case of a asymmetric traffic profile.
IP | Index Pulse – pulse signal triggered by an indicator to define a reference point for each rotation
IP | The IP Code, International Protection Marking, IEC standard 60529, classifies and rates the degree of protection provided against intrusion
IP | Internet Protocol
IPC | InterProcessCommunication
IP-core | reusable part of the FPGA code within the SiDaNet device module in terms of a verified net list
IPMI | Intelligent Platform Management Interface (supports remote control functionality)
IPS | Image Processing System
IPT | Image Processing Toolkit
IQ | Image Quality
IR | Iterative Reconstruction (rawdate-based)
IR | Interventional Radiology
IRG | Interactive Region Growing
IRIS | iterative reconstruction in image space
iRMA | "integrated Requirements Management Tool" This Project covers the customization of the purchased tool CaliberRM for H IM CR.
IRS | Image Reconstruction System
IS | Information System
IS | Information Security
ISD | Inter Segment Dependency
ISDN | Integrated Services Digital Network
ISM | Industrial, Scientific and Medical Band
ISO | International Organization for Standardization
IT | Integration Test
IT | Information Technology
IT Administrator | The IT Administrator is the administrator of e.g. a hospital and is different from the clinical user and has only some features regarding administering the system.
IT_PD | IT-Power Distributor
ITE | Image Type Exchange
ITE | Information Technology Equipment
ITR | Iterative Reconstruction
ITSP | Integration Test Specification
ITV | Internal Target Volume
IVB | Image Volume Base
IVJ | Intervention Joystick
IVK | Installed Volume Components
IVK / IVK + | IVK Installierte Volumenkomponente (TD Kürzel f. LINA-pflichtige Teile) / Installed Volume Components
IVM | Interventional Module
IVP | InterVention Panel
IWP | Imaging Work Place
Jacket | charger and holding device for jacket (with tablet-PC)
JPG | Graphics file type/extension
JTAG | Boundary scan and FW load bus
K | Kelvin
KERMA | Kinetic Energy Released per unit Mass of Absorbe
Key Image Note | The Key Image Note Integration Profile specifies a transaction that enables a user to flag as significant one or more images in a study by referencing them in a note linked with the study. This note includes a title stating the purpose of the flagged images and a user comment field. These notes will be properly stored, archived and displayed as the images move among systems that support the profile. Physicians may attach key image notes to images for a variety of purposes: referring physician access, teaching files selection, consultation with other departments, and image quality issues, to name a few.
Keys | Control Box and Gantry: - START key - STOP key - SUSPEND key - MOVE key - ZERO POS key Keyboard
KFDA | Korea Food and Drug Administration
KFDA | This integration profile includes both the department imaging systems and the hospital-wide image distribution such as:
kg | Kilogram
KGW | Knowledge Gateway
KIN | IHE Integration Profile Key Image Notes
kPa | kilopascal
KPI | Key Performance Indicator
kV | Kilovolt
KVM | Keyboard Video Mouse
KVM Switch | Keyboard Video Mouse Switch
KVM-over-IP | Keyboard Video Mouse over Internet Protocol
kVp | Kilovolt Peak
L2V | Lumen to Vessel
LA | Left Atrium
LAD | Left Anterior Descending
LAN | Local Area Network
LAO | Left Anterior Oblique
LAP | Laser systems for measurements and projection
LASER | Light Amplification by Stimulated Emission Radiation
LAT | lateral
LAV | Low Attenuation Volume
LB | Large Bore
LB-HP2 | P83A (large bore, HP2)
LB-HP4 | P83B (large bore, HP4)
LCA | Left Coronary Artery
LCB | Line Connection Box
LCE | Low contrast edge-enhancing filter
LE | Less than or Equal to
LED | Light emitting Diode
LEM | Layout Editing Mode
Leonardo | Siemens workstation product
LHem | Left Hemisphere
LHS | 'Left Hand Side' of the AQWP displaying on the left monitor the user interface of the scan and reconstruction software that is running on the ICS.
Life | Siemens program concerning customer relationship, see also WorldCare Customer Retention Program (CRM) for Siemens CT scanners.
LIMA | left internal mammary artery
LINAC | Linear Accelerator
Link | A wireless, wire or fiber point-to-point connection between two SiDaNet ports. Each links establishes a bi-directional, full-duplex connection. A bandwidth of a bi-directional link could be increased by a additional unidirectional link in the case of a asymmetric traffic profile.
LINKS | stored association between ELEMENTS and external sorces
LISA | In-Room Scan Application
LISN | Line Impedance Stabilisation Network
List | Locally one or multiple Work Items and/or Test Results. Those are read from TFS WorkItemStore (VSTS) and Microsoft Test Case Management Component. Each List can be exported to a defined Format.
Live image / LRS | Real time image (known as RTD)
live media | Media, e.g.DVD, that allows using applications/data files from media without installation
Live Reconstruction / Live Recon | The Live Reconstruction Series is the result of an Online Reconstruction to provide immediate visual feedback during scanning by a live visualization of reconstructed slices which respectively fit to the most recent or current scan position. Use cases for that functionality are the observation of contrast media flow, interventional scanning (e.g. "CT iFluro") or simply providing perceivable scan progress to the user. The parameterization of a Live Reconstruction Series is optimized for least reconstruction time and not best image qualtiy. Hint: Live Reconstruction Series corresponds to the term Real Time Display Reconstructions/RTD.
LL | Lower Left
LLL | Left Lower Lobe
LLS | Lung Lesion Segmentation
LM | Light Marker
LM | Left Main
LMAS | "Liegen" Master on PHS
LMB | Left Mouse Button
LO | Laser Origin
Loaded image | Any of the images loaded into the Dynamic Evaluation application. These are the images used as or used for computing the baseline image and the dynamic images.
loading | irradiation
Log repository | Its purpose is to store the Log Messages. Repository can be etc. a file, the Windows Event Log or a DataBase.
Logical Slice Number | Simply counts the transmitted slices within a reading from 1 to<n>, where n depends on the number of selected slices. The header always has a slice number of zero (physical and logical).
LOID | Logical Object Identifier
LP | Line Pairs
LP | Low Power
LP1 | P81A (low power, 1 cm)
LP2 | P81B (low power, 2 cm)
LPDDR | Low Power Double Data Rate memory
LPS | Light Position Sensor
LR | Lower right
LRS | Live Reconstruction Series
LS | Loudspeaker
LSB | Least Significant Byte
LSB | Least Significant Bit
LSD | Link Speed Divider, divider between SSS and CCS who passes only every n Master Transmit Moment signal CSI_t_Tmtx.
LTDA | Long Term Data Archiving
LUL | Left Upper Lobe
LungRADs | Lung Imaging Reporting and Data System
LUT | Lookup Table
LV | Left Ventricle
LVA | Left Ventricle Analysis
LVD | Low Voltage Directive
LVDS | Low Voltage Differential Signaling
LWL | Lichtwellenleiter
M1 | Market segment no. 1. Market segmentation in a range from 1='high' to 4='very low' mainly concerning the product pricing.
M2 | Market segment no. 2. Market segmentation in a range from 1='high' to 4='very low' mainly concerning the product pricing.
M3 | Market segment no. 3. Market segmentation in a range from 1='high' to 4='very low' mainly concerning the product pricing.
M4 | Market segment no. 4. Market segmentation in a range from 1='high' to 4='very low' mainly concerning the product pricing.
M6T | Market segment 6 with X-ray tube Vectron (Taifun)
mA | Milliampere
mA s | Milliampere Second
MAB | Medical Advisory Board
MAC | Media-Access-Control
macro | A macro is a set of functions for creating or modifying formattings in MS Word documents.
MAF | Memory Access Function, controls the memory access inside the DAS.
Mains Failure | line input to system doesn't provide power within the specified range, e.g. no power is supplied
MAR | MAster Rotating
MARS | Molecular Acquisition & Reconstruction System
mAs | Milliampere second
MAS | MAster Stationary
mb | mouse button
MB | Megabyte (10^6 Bytes)
MB/s | Megabyte (10^6 Byte) per second
MBC | MonoBlock Controller
MBF | Multiband Filter
MBF | Myocardial Blood Flow
MBH | Modified beam hardening correction
MBI | Memory Bus Interface, allows the connection of a dual port RAM to handle the memory data type. The MBI consists of address- and data bus and control signals.
MBit | 10^6 Bits
MBV | Myocardial Blood Volume
MC | Motion Correction
MC | Motor Controller
MCE | Master Control Extension
MCS | Monitor Ceiling System
MCS | Modulation and Coding Scheme (MCS) index value defined in the IEEE 802.11 standard.
MCS 1 | Monitor Ceiling System 1
MCS 2 | Monitor Ceiling System 2
MCS pro | Monitor Ceiling System professional
MCSF | Memory Confirmation Subfunction
MCSP | Mean Costs Spare Parts
MCSPCpS | Mean Costs Spare Parts consumed per System and Month sum of spare part costs (*) / number of systems notification categories "repair totally down", "repair partially down" and "scheduled reactive service", without expendable parts (*) 43.4% of customer
MCTR | Mean Confirmed Time Repair = sum of confirmed duration of work / number of notifications with confirmed data
MCWH | Mean Confirmed Working Hours per System and Month = sum of confirmed duration of work / number of systems
MD | Measuring Device
MDD | Medical Device Directive
MDM | Measurement Data Module
MDS | Maximum Dose Saving
MDS2 | Manufacturer Disclosure Statement for Medical Device Security
MDT | Mean Downtime = (downtime 1 + downtime 2) / number of notifications with site visit
ME | Myocardial Enhancement
ME | Mechanical engineering
ME | Myocard Enhancement
ME | medical electrical
ME EQUIPMENT | Medical equipment
ME SYSTEMS | Medical Equipment Systems
MEASURE M | means to reduce the risk
Measurement | A measurement is the smallest entity that can be triggered for execution by user. Example could be an Auto Range or a single Range (e.g. a "spiral" acquisition).
MED QR | Siemens medical Quality Requirements
MedBook | Templates for official documents at Siemens Medical, containing macros and applications.
MedCom | Former name of syngo. Medical Common Software (incorporates former CSA and CAP)
MEDHOME | Environment variable containing base path of syngo system, e.g. of CT system C:\Somaris or of MMWP C:\Leonardo -- use as follows: %MEDHOME%
MedTrace | Trace analysis tool
MEF | Managed Extensibility Framework
Memory | Data type which allows a remote read and write access to the memory of a other device. The memory is organized in linear addresses, the address space is 64 Kwords with 16 bit each.
MEP | Market Entrance Phase (project phase)
Message Serverity | the message handling provides messages with the severity Information, Warning, Error
META-DATA | META-DATA are the data which describe a document (document information): e.g. - Productname - Productversion - Organisation - Title of document - Status of document - Author - Index - U-Stand ...
METRO | This is a functionality parallel to scanning and reconstruction.
METRO 3D | Reconstruction of instant 3D images, see Instant 3D
METRO Archive | Archival of images after the end of the series as a background function.
METRO Copy | Network transfer of the reconstructed image to a DICOM node parallel to other activities as soon an image is reconstructed.
METRO Delete | Delete patients, images, raw data in the background, following defined rules.
METRO Filming | Filming as background function
METRO Recon | Reconstruction of images in parallel but asynchronous to data acquisition.
MEZ | Mehrfachfehler Eintritts Zeit, Multiple fault occurrence time
MFC | Microsoft foundation classes
MFM | Multiport Fastlink Master
MFOT | Multi-Fault Occurrence Time
MFR | MED Fachliche Richtlinien
MFR | Med Technical Guidelines (MFR)
mGy | Milligray
MHU | Mega Heat Unit
MI | Molecular Imaging
MI.One | Software for hybrid PET-CT or SPECT-CT systems
MiB | Mebibytes (2^20 Bytes)
MIB | Management Information Base (describes information that can be obtained and/or modified via a network management protocol, e.g. SNMP)
MIC | MICrofon
MICa | Microphone cabinet: A small cabinet with mounting facilities for MEMS microphone(s)
MID | Module Interface Demultiplexer (IP-core)
MIMO | Multiple Input Multiple Output
MinIP | Minimum Intensity Projection
MIP | Maximum Intensity Projection
MIP image | Maximum Intensity Projection image. A postprocessed image of the same size as the input images where the value at each pixel (xi, yi) is the maximum enhancement value reached in any one of the input images at that location. The pixel values are in HU.
MIRSF | Memory Indication and Response Subfunction handles the indication and response primitives in the MAF.
MITA | Medical Imaging & Technology Alliance
MIT-BIH | Massachusetts Institute of Technology-Beth Israel Hospital
ML | Middle left
MLC | Multi Leaf Collimator
MLD | Mean Lung Density
MLJ | Marketing License Japan
MM | Multimodality
MMCA | Multi Modality Cardiac Application, "predecessor" to MI in Circulation
MMIO | Memory Mapped IO
MMWP | Mulltimodality Workplace
MNP | Managed Node Package
MOD | Magnetic Optical Disc
Modality Worklist | A Modality Worklist is a list of Scheduled Procedure Steps. It captures all the attributes from the IS, that are mandatory to be inserted into the resulting DICOM Image Object(s).
Mode entry | one mode loaded over the SCI
MOMO | Message Observer Monitor
MOOP | Means of Operator Protection
MOPP | Means of Patient Protection
MOTS | Modifiable off-the-shelf
MP | Master Page
MPG | Medizinproduktegesetz (MDD)
MPN | Manufacturer Part Number
MPPS | Modality Performed Procedure Step (DICOM Service)
MPPS | Modality Performed Procedure Step
MPR | Multi Planar Reconstruction
MPT | Medium Power Transmission
MPT | Multi Purpose Table
MPT2 | Multi Purpose Table 2
MPVV | Multi Phase Volume Viewer
MQR | Med Quality Regulation
MR | Magnetical Resonance
MR | Middle right
MRI | Magnetic Resonance Imaging
MRQSF | Memory Request Subfunction, handles the request primitive in the MAF.
MRS | Market Requirement Specification
MRT | Mean reaction time; calculated as MRT [h] = Reaction time [h] / Number of service actions
MRTC | Mean Repair Time (confirmed) sum of confirmed duration of work [hours] / number of notifications with confirmed data [notification categories "repair totally down", "repair partially down" and "scheduled reactive service"]
MRTT | Mean Repair Time Task
MS | Microsoft
ms | Millisecond
MS | Millisecond
MSB | Most Significant Bit
MSB | Most Significant Byte
MSI | Multiplexed Signal Interface, contains the signals in the same (multiplexed) format like inside the cells.
MSI | Microsoft Windows Installer
MSP | Microsoft Windows Installer Patch
MSPC | Mean Spare Part Costs per System and Month = sum of spare part costs / number of systems
MST | Microsoft Transform
mSv | Millisievert
MTBF | Mean Time between Failure
MTBFwSPC | Mean Time between Failure with Spare Part Consumption
MTBFwSV | Mean Time between Failure with Site Visit
MTBN | Mean Time between Notifications operating months / number of notifications notification categories "repair totally down", "repair partially down", "scheduled reactive service", "maintenance" and "update"
MTBSC | Mean Time between Service Call
MTF | Modulation Transfer Function
MTM | Microsoft Test Manager
MTM | Master Transmit Moment, defines the transmit moment inside each component. All ports of a component start the transmission of the cells synchronous.
MTP | Master Timing Port, network of a switch in SyncRepeater configuration, from which her derives his synchronization information.,
MTT | Mean Transit Time
MTTF | Mean Time to Failure
MTTR | Mean Time to Repair
Mul | Multiscan
Multi Study Examination | A Multi Study Examination is an examination with more than one Requested Procedure. This name is chosen since each Requested Procedure corresponds to a seperate Study.
multimodality system | a combination of a CT with another modality e.g. PET-CT or SPECT-CT
multiple tomogram system | multislice CT
Multirow | Number of detector elements in z-direction > 1
Multiscan | A scan where Xray is on for multiple turns of the measurement system (spiral scan without table feed) to acquire data of the same anatomical region over time, used for perfusion measurement or contrast flow studies
Multislice | Number of acquired slices with one rotation > 1
MV45 | microphone amplifier for P45
MVC | Model View Controller
MVT | Medium voltage transformer
MVVM | Model View ViewModel
MWL | Modality Worklist
MxT | A substitute for M9T and M6T which are the synonyms of the new CT scanners
MxT | A substitute for M9T which is the synonyms of the CT scanner
n | number counter
N | Newton
n.a. | not applicable
n/s | native/syngo - Software platform from business line Syngo
NAD | no abnormality detected
nAIF | normalized Arterial Input Function
NAP | Needle Artefact Prevention
NAPT | Network Address Port Translation
NAT | Network Address Translation
Native scan range | A scan range without a Contrast protocol
Navigator | marketing name for ICS (shall not longer be used as abbreviation)
NCB | National Certified Body
NCM | Non-conforming Material
NEC | National Electric Code
NECT | Non enhanced CT
NEMA | National Electrical Manufacturers Association
Neuro DSA CT | neuro digital subtraction angiography CT
NEV | Nodul Enhanced Viewing
NFPA | National Fire Protection Association
NGFF | Next Generation Form Factor
NHIN | National Health Information Network
NIBP | Non-invasive blood pressure
NLS | Native Language Support
NMPA | National Medical Products Administration is the Chinese agency for regulating drugs and medical devices (formerly the China Food and Drug Administration or CFDA)
NMT | CANopen Network Management
NOC | No Objection Certificate
non-CBR | Reconstruction without cone-beam correction
Non-Content Work Items | All Work Items stored in the database that are required to manage and support the handling of Content Work Items (e.g. Baseline, Test Review, ...)
nP | non-Product
NPU | Network Processing Unit
NPU | Network Processing Unit: electronic board which contains SiDaNet Switch, unit to control CDT modules and extracts Measurement data stream to external ethernet interface
NPU DM | NPU device module: variant of the NPU that provides SiDaNet device module instead of SiDaNet switch functionality.
NRTL | Nationally Recognized Testing Laboratory
NRTL | Nationally Recognized Testing Laboratory (USA)
NRU | Network Routing Unit
NSA | National Security Agency
NSS | National Security Systems
NTFS | New Technology File System
NTP | Network Time Protocol is a networking protocol for clock synchronization between computer systems
NTP | Account for H IM products
NVMe | Non Volatile Memory express
nVolume Compare | Clinical Task provided by SY which allows the user to compare multiple volumes in an easy way.
OAC | Online Adaptive Control
OAR | Organ at Risk
OC | Open Collector
OCT | Operative Clinical Task (A task of clinical relevance being part of a taskflow, e.g. reading. An OCT can technically be realized using one or more technical tasks.)
OCX | OLE custom control
OEM | Original Equipment Manufacturer
Offline Reconstruction / Offline Recon | Scanning, respectively performing a Range usually will acquire Raw Data. In contrast to an Online Reconstrustion where the reconstrcution of that data is performed during the scan process, an Offline Reconstruction can be performed at any time after the Range execution has finished. The purpose of Offline Reconstrustions is to produce the requested diagnostic image data output that is expected from the examination. Offline Reconstructions are therefore usually predefined steps of the examination protocol. They can be configured to start automatically after scanning, but also can be manually added, changed regarding parameterization and triggered by the user. Depending on the actual type of the examination there might be also different kinds of an Offline Reconstrcution regarding user interface, possible parameterization and reconstruction algorithms (e.g. cardio phase based reconstructions).
OGC | Optical Guidance Control
OIS | Oncology Information System
OLH | Online Help
OM | Operator Manual (legal required documentation)
Online Reconstruction / Online Recon | The purpose of an Online Reconstruction is to keep up with the scan of a Range in the sense of immediately reconstructing the currently acquired Raw Data. The term might be translated into “Reconstruction while the scan is online”. In terms of time it means that both, scanning the Range and Online Reconstruction are performed close together, whereas the latter one usually, but not necessarily finishes immediately after the acquisition of data has ended. The result of an Online Reconstruction is an Online Reconstruction Series. In the current version Online Reconstruction Series are Live Reconstruction Series and Monitoring Series.
OpenGL | Open graphics library
Operating Volume | The operating volume is defined where the physician can effectively perform intervention with the laser needle guidance system. The needle entry point should lie within this volume.
operational state | state, in which PatPos movements and X-ray is possible. Also called: ‘closed’ safety circuit. Opposite: safe state or ‘open’ safety circuit
OPM | Output Position Multiplexer, second multiplexer hierarchy inside the SMM.
option | Functionality which can be bought in addition to the standard configuration, usually protected by one license key.
ORA | Optimized Reconstruction Algorithm
OS | Operating System
OSB | oriented strand boards
OSC | Operation System Interface and Communication
OSHA | Occupational Safety and Health Administration (USA)
OSHPD | Office of Statewide Health Planning and Development
Osteo CT | Bone mineral density measurement and semiautomatic evaluation method for the quantitative assessment of bone mineral content in the lumbar spine.
OT | Other; used for dicom objects which do not belong to a modality
OTS | off-the-shelf
OVP | Over Voltage Protection
owner | The owner is responsible for the definition and fullfillment of the requirement
P06 | m1/m2 CT platform
P08 | m3 CT platform
P45 | CT Project 45
P46 | CT Project 46
P47 | CT Project 47
P58A | SOMATOM Force
P60 | CT Project 60
P60A | CT System SOMATOM X.cite
P60B | CT System P60B
P63A | CT System P63A
P70 | CT/D2I Project 70
P70A | CT/D2I System P70A
P70A | CT System SOMATOM On.site
P79MA | P79 Mobile Application
P81 | CT Project 81
P81A | CT System SOMATOM go.Now
P81B | CT System SOMATOM go.Up
P82 | CT Project 82
P82A | CT System SOMATOM go.All
P82B | CT System SOMATOM go.Top
P82C | CT System SOMATOM go.Fit
P83 | CT Project 83
P83A | CT System SOMATOM go.Sim
P83B | CT System SOMATOM go.Open Pro
P2408 PET | CT Project 2408
PA | Posterior-Anterior
Packet | Data type, which allows the transportation of data fields with up to 64Kbytes via SiDaNet. The data field is transported with more than one cell, the segmented information elements of each cell are reassembled at the destination to the original packet.
PACS | Picture Archiving and Communication System
Page Mode | In page mode filming, the entire film sheet is built up in the computer and then transferred to the camera.
PAI | Packet Application Interface, connects the cells, which are part of a packet, to the application which performs the segmentation and reassembly (SAR) process.
PAL | Programmable Array Logic
Parameter image | An image created by postprocessing operations where the pixel values are one of the parameters computed from the set of input images, e.g., Time-to-Peak image, Peak-Enhancement image.
Parameter List View | The UI control that represents the set of range, recon and injection phase parameters of the current examination protocol.
Parameter Panel | The Parameter Panel is a UI control which allows the user to view and edit parameters for scanning and reconstruction.
PAT | Port and Address Translation
Patient | A Patient is a person receiving, or registered to receive, healthcare services, or is the subject of one or more studies for some other purpose, such as research (see Digital Imaging and Communications in Medicine (DICOM), NEMA, 2013). In context of Somaris/10: Patient is a Clinical Patient or a Service Patient (e.g. phantom). Clinical Patient is a Scan Patient or a Recon Patient. Scan Patient is a patient selected for scanning. Recon Patient is a patient selected for reconstruction.
Patient | represents the patient that is examined in a workflow
Patient data | Patient data include demographic data, reports, lab results, but no image data
Patient Information Reconciliation | This Integration Profile extends Scheduled Workflow by providing the means to match images acquired for an unidentified patient (for example, during a trauma case) with the patient’s registration and order history. In the example of the trauma case, this
PatPos movement | all movements of the patient table relative to the gantry, which are relevant for safety, i.e. horizontal, vertical and tilt movement
PB | Patient Browser
PBU | Parallel Backprojection Unit
PBV | Perfused Blood Volume
PBX-10 D6 | Parallel Backprojector Board, partnumber 83 69 998
PBX-40 D5 | Parallel Backprojector Board, partnumber 77 15 969
PBX-60 D7 | Parallel Backprojector Board, partnumber 83 70 004
PC | Personal Computer
PC_Target | System with non syngo platform
PCAP | Projected Capacitive Touchscreen
PCB | Printed Circuit Board
PCBA | Printed Circuit Board Assembly
PCCT | Photon Counting CT
PCD | Photon Counting Detector
PCI | Peripheral Component Interconnect
PCIe | Peripheral Component Interconnect Express
PCI-X | Peripheral Component Interconnect - extended
PCON | Power Controller (SPECT)
PCS | Packet Control Structure, contains the information, how much real data bytes are transported inside the PTU. Additional a packet counter and a user sharable byte are included in the PCS.
PDB | Power Distribution Box
PDB | Power Distribution Board
PDC | Power Distribution Cabinet
PDF | Portable Document Format
PDF/A | PDF/A is an ISO-standardized version of the Portable Document Format (PDF) specialized for the digital preservation of electronic documents. PDF/A differs from PDF by prohibiting features ill-suited to long-term archiving, such as font linking (as opposed to font embedding).
PDI | IHE Integration Profile Portable Data for Imaging
PDO | CANopen Process Data Object
PDO | P-path Digital Output (not self-tested)
PDR | Power Distribution Rotating
PDS | Power Distribution System
PDS | Power Distribution Stationary
PDS | Prospective Dose Saving
PDU | Power Distribution Unit
PE | Pulmonary Embolism
PE | Protective Earth
Peak-Enhancement | Peak enhancement in HU at a pixel location (xi, yi) is defined as the difference in maximum enhancement of the pixel value reached in any one of the input images at the given pixel location (xi, yi) and the enhancement in the baseline image at that location.
Peak-Enhancement image | A postprocessed image of the same size as the input images where the value at each pixel (xi, yi) is equal to the Peak-Enhancement value (within the input images) at that location. "Peak-Enhancement image = MIP image - Baseline image"
PEMS | Programmable Electrical Medical System
PEP | Produktentstehungsprozeß (product development process)
Permissions | Authorization to access specific data using a well-defined access mode ("create, read, update, delete, execute"). For functional security, only the "execute" flag is used. For data security, all five flags are used, where "execute" has the meaning to provide authorization to change protections on a data object.
PESS | Programmable Electronic Subsystem
PET | Positron Emission Tomography
PEV | Polyp Enhanced Viewing
PFC | Power Factor Correction
PFL | Panel at the Front side Left
PFO | Posterior Fossa Optimization (Filter algorithm to reduce beam hardening artefacts)
PFR | Panel at the Front side Right
PGC | PET Gantry Controller with SiDaNet
PGP | Presentation of Grouped Procedures (IHE profile)
PHS | Patient Handling System
PHS4 | Patient Handling System 4
PHY | Physical Layer
PI | Power Injector
PIB | General Purpose Interface Board
PIR | Patient Information Reconciliation (IHE profile)
Pitch | Term applicable only to SPIRAL CT. It is defined as the ratio of table feed per rotation to single slice collimation.
Pitch | Ratio of table feed per rotation and beam collimation: pichfactor = f / (n*w) with f = table feed per 360° rotation: n = number of collimated slices; w = collimated slice width
Pixel | Picture Element
Pixel value | Value of picture element, for CT described in Hounsfield value.
PK | Prüf Kabine
PKI | Public-Key-Infrastructure
Planning Base | A Planning Base is an image Series used for planning of Reconstructions. There are two different types of Planning Bases: - Frame Set Planning Base and - Volume Planning Base. In CT environment all Reconstruction Series are of type Volume Planning Base except for multi-phase reconstructions which are covering the same volume multiple times.
PLD | Programmable Logic Device
PLI | Physical Layer Interface, media independent interface to connect the data link layer to the physical layer.
PLL | Phase Locked Loop
PLV | Paramter List View
PM | Protection Module
PM | Protection Module: safety component of a CT system
PMC | Patient Motion Controller (PET)
PMIC | Power Management Integrated Circuit
PMM | Physiological Measurement Module
PMMA | polymethylmethacrylate; acrylic phantom material
PMT | Photo Multiplier Tube
PN 15 | Pseudo random Noise, a sequence with2^15-1 length
PoE | Power over Ethernet
POF | primary operating function
POI | Point of Interest
Port | Connection between a SiDaNet component and a SiDaNet link.
Port | is an endpoint of communication in an operating system
PostProcessing Workflow | IHE Integration Profile with the objectives of - Schedule and track post-processing workflow steps using standardized mechanisms - Share workload among workstations, chain workflow steps - Performed work status sharing - Store and Query/Retrieve of Evide
Power Cycle | describes a situation when power is disconnected from a subsystem or the whole system and reconnected after a certain time, e.g. cold start, mains failure and power return
Power Save | if the system is switched on, but currently not in use, power consuming components get allowed to enter less power consuming states or get switched off and reactivates automatically on user interaction
Power States | during operation eSW has to switch the system/components into different configurations of power, named Power States
PP | Parameter Panel
PP | PlatformPart
P-path | Protection Path
PPM | Parts Per Million
PPM | Patient Position Monitor (SPECT)
PPM | Performance PM
PPT | Pseudo Position Trigger
PRBS | Pseudo-Random Binary Pattern
PRBS7 | Pseudo-Random Binary Pattern with length 2^7-1
Precision | Difference between several independent measurements taken under the same conditions.
PREP | Preprocessed Data
PREP | Product-related environmental protection = Produktbezogener Umweltschutz
Presentation of Grouped Procedures | The PGP Integration Profile addresses the complex information management problems entailed when information for multiple procedures is obtained in a single acquisition step (for example CT of the chest, abdomen and pelvis).PGP provides the ability to view
PRF | Variable Transmit Function, controls the handling of a received cell with packet data inside the DAS.
Print Composer | A system that generates DICOM print requests to the Print Server. Print requests include presentation state information in the form of Presentation Look-Up Tables (Presentation LUTs).
Privilege | Authorization to perform a specific functionality
PRL | Panel at the Rear side Left
Proactive Service | service interaction that corrects potential malfunction before system operation is limited or restricted to customer
PROBABILITY | estimation of the potential occurrence of harm
Prognostic Data | Any information of a component or system that allows proactive service or any information of a component or system that indicates a malfunction of a component in the near future
Projection | All readings with different focal spot positions belonging to the same tube angle
PROJECTs | collection of multiple subsets of ELEMENTs
ProMED tool | a Microsoft Excel application to assist in obtaining, recording and managing product ingredients. The data may be stored within SAP
Protections | Security token for database objects
Protective Function | abstract measure intended to overcome the risk of one or several Safety Function(s), e.g. “protection against unintended motion”
Protocol | The Protocol is sum of all steps performed during an Examination or Service Application that require either an interaction by the user or are performed automatically by the system. Examples for these steps are Measurements. some user instructions or some automatic checks.
protocol element | a set consisting of all the scanning parameters necessary to perform a scan
PRR | Panel at the Rear side Right
PRS | Platform Requirement Specification
PSC | Power and Safety Controller
PSD | Position Sensitive Device
PSR | Power Slip Ring
PSSO | Product Security Solution Officer
PSSO | Product and Solution Security Officer
PT | Particle Therapy
PTAB | Patient TABle
PTAB | firmware running on patient table control board
PTC | Positive Temperature Coefficient
PTF | Packet Transmit Function, controls the generation of a cell with packet data for transmission inside the DAS.
PTO | Port Transmit FIFO, transmit FIFO inside the SAF for each timeslot.
PTU | Packet Transmission Unit, total number of bytes required to transport a packet via SiDaNet. This includes the real data bytes, idle bytes and the Packet Control Structure.
PTV | Planning Target Volume
PTZ | Prozessfehler Toleranz Zeit, Process fault tolerance time
PUBLISH | export data in defined file-formats with a defined layout structure and content preperation
PUD | Power Unit Detector
PULMO | Name for the program that is described within this document
Pulmo 3D | Marketing name for Chest Reading task
Pulmo CT | Evaluation of the CT values of the lung tissue.
put on the market | when product leaves the manufacturing facility (same country) or is imported into the country where it is to be sold with the date either being the date a product ships from a manufacturing facility (when shipped within a country) or the date it is imported to the country where the product is sold to the end user.
PV | Protocol Version, identifier to differ between the current a newer protocol version of SiDaNet.
PVE | Packaging label (German: Produktverpackungsetikett)
PVP | Portal Venous Perfusion
PWB | Printed Wiring Board
PWF | Postprocessing Workflow (IHE profile)
PWM | Pulse Width Modulation
PWM | Perfusion Weighted Map = simple transformation of a CTA series by segmentation, filtering and colour mapping; well-suited for 3D stroke reading
PXR | Proxy X-ray System
PXR | Power & X-ray
Q | Quality
QA | Quality Assurance
QAR | Quickscan Artifact Reduction
QC | Quality Check
QCA | Quantitative Coronary Analysis
QGS | Quantitative Gated SPECT The Cedars Sinai Quantitative Gated SPECT (QGS) option enables automated processing of myocardial perfusion SPECT and gated SPECT data. Provides left ventricular ejection fraction calculations, 3D beating image displays, volume curves and polar maps.
QM | Quality Management
QM | Qualitiy Management
QMP | Quality Management Plan
QoS | Quality of Service, network guarantees agreed limits for bandwidth and delay
QPCA | Quantitative PET Cardiac Analysis program
QPS | Quantitative Perfusion SPECT The Cedars Sinai Quantitative Perfusion SPECT(QPS) package is an interactive standalone application for the automatic segmentation, quantification, analysis and display of static (ungated) short axis myocardial perfusion SPECT data.
QR Med SAG | Quality Requirements Med SAG
QRS | Quasi-Random Signal
QRS | sequence of characteristic waves within an ECG
QSR | Quality System Regulation (formerly GMP: Good Manufacturing Practices)
QVA | Quantitative Vascular Analysis
R | Review Marker
R&D | Research and Development
R_MOT | Rotation MOTor
RA | Risk Analysis
RAC | Rotating Anode Control
Radio | Radiographic image, special reconstruction like conventional x-ray.
Radiologist | A radiologist is a clinical admin.
RADS | Reporting and Data System
RAID | Redundant Array of Independent Discs
RAL | color codes
RAM | Random Access Memory
Range | A Measurement contains 1 to n Range(s). A Range specifies a certain type of data acquisition and the area where that data shall be acquired. Furthermore it represents an abstraction that shall hide further execution specifics of the respective acquisition type. It therefore holds 1 to n Scans. E.g. a "Spiral" Range contains a single Scan, whereas a "Sequence" Range may contain more than one Scan.
Range Task | Task in the Taskflow Control representing a Range.
Ranorex | A test automation framework for automated testing of desktop, web and mobile applications
RAO | Right Anterior Oblique
raw data | Raw data are available after preprocessing. Raw data means data which can be used to calculate images within a certain range of slice thickness, FoV and any available kernel.
RC | Remote Control (CT)
RCA | Right Coronary Artery
RCD | Residual current device
RCM | Residual Current Monitor
RCS | Reference Coordinate System
RD | Remote Diagnostics
RDB | Reading Data Buffer
RDC | Remote Desktop Connection
RDIAG | Remote DIAGnostics Server, intranet portal for service engineer to connect to customer CT console via SRS
RDP | Remote Desktop Protocol is a proprietary protocol developed by Microsoft, which provides a user with a graphical interface to connect to another computer over a network connection.
RDS | Remote Diagnostic Service
RDS - SRS | Remote Diagnostic Service [Siemens Remote Services] is an emerging speciality and it offers many advantages over the conventional service facilities. The connection between the CT system and the remote diagnostic station (SRS server) is etablished via a dedicated telefon line (analog or ISDN router, modem) or via the Internet (VPN tunnel).
RE | Radiated Emission
REACh | Registration, Evaluation and Authorization of Chemicals
Reading | Data of all detector channels and slices generated at one focal spot position
Reading Data Structure | Is a format for a set of data containing control and measurement information for a single detector position (data of multiple slices for a multi slice scanner). The Image Reconstruction System (IRS) expects the received data to be compliant to this format. When received via the serial input the same data structure is used, but with transmission protocol overhead - control words and received DAS-CRC.
Reading Number | Each received block contains a 32 Bit number, incremented with each reading within a scan.
ready state | preparation of the device
Real Time Display | Displaying of the measured image synchronous to the scanning and just after the aquisition of a slice (rotation) is finished.
Real Time Recon | Reconstruction synchronous to scanning with an additional delay for displaying the first image.
REC | RECeiver
Receiver-CRC | CRC calculated for the contents of a block excluding the footer for transmission from Receiver to IRS
RECIST | Response Evaluation Criteria in Solid Tumors (oncology review criteria)
Recon Duration | The time elapsed from starting to finishing a recon job on the IRS
Recon Job | A Recon Job is the activity to perform a Reconstruction run in the background.
Reconstruction Planning Base | Synonymous to Planning Base
Reconstruction time | Time between the start of reconstruction until the image of that scan appears on the monitor. It is presumed, that the reconstruction starts after 180° of data collection.
RED | Radio Equipment Directive
Reference Image | reference image within the graphical 3D recon planning (GRP) to show the orientation of an oblique image
Reference Lines | topogram image with graphic reference to scans or images (topographics)
RefV | Reference Vessel
RELATION | stored association between ELEMENTS
RELATIONTYPE | abstract definition of possible association between ELEMENTTYPES
release note | information about content of software update and last minute changes
Remote Application | A Remote Application Server is the server that provides all Remote Applications to users through the Remote Desktop Protocol (RDP)
Remote Desktop Services | Remote Desktop Services is one of the components of Microsoft Windows (both server and client versions) that allows a user to access applications and data on a remote computer over a network, using the Remote Desktop Protocol (RDP).
Remote Diagnostic Service | Remote Diagnostic Service is an emerging speciality and it offers many advantages over the conventional service facilities. RDS needs a modem HW at the CT site, dedicated telephone or ISDN connection between CT site and the remote diagnostic station and R
RemoteApp | RemoteApp is a special mode of Remote Desktop Services where the remote session configuration is integrated into the client operating-system. In RemoteApp mode, only the content of a single application is rendered on the client desktop behaving like any other client window.
REMOVE | mark information as no more available in the system
REP | Reportsheet
REPO | Removal Emergency Power Off
Report Creator | A system that generates and transmits draft (and optionally, final) reports, presenting them as DICOM Structured Reporting Objects. It may also retrieve worklist entries for reporting steps from the Report Manager and provide notification of completion of the step, allowing the enterprise to track the status of an awaited report.
Request | First step in a communication interaction where the data transmission is invoked at the source device.
Requested Procedure | A Requested Procedure is an instance of a Procedure of a given Procedure Type, which is an item in a catalouge of imaging procedures that can be requested and reported upon in an imaging service facility. A single Requested Procedure is the smallest unit of service that can be requested, reported, coded and billed. A Requested Procedure leads to one or more Scheduled Procedure Steps involving Protocols as specified by a Procedure Plan (see 7.3.1.5 in PS 3.3, Digital Imaging and Communications in Medicine (DICOM), NEMA, 2013).
requirement | The IEEE Standard Glossary of Software Engineering Terminology (1997) defines a requirement as: (1) A condition or capability needed by a user to solve a problem or achieve an objective. (2) A condition or capability that must be met or possessed by a system or system component to satisfy a contract, standard, specification, or other formally imposed document. (3) A documented representation of a condition or capability as in 1 or 2.
requirement grid | The Requirement Grid is used to display a set of requirements. The grid view includes e.g. the requirement type, tag, name, status, version, priority, owner and many other attributes. The user can also select a requirement and view its description in the lower pane of the Requirement Grid. Furthermore the user can sort and filter the requirements in the Grid.
requirement key | A requirement key is a unique key with a defined syntactical structure used to identify requirements
Response | Mechanism to delay the feedback to the source device during a bi-directional communication procedure. When the feedback could not be sent inside the confirmation function, the response function could be called later when the feedback information is available.
Result image | Any of the images created by the Dynamic Evaluation application (Baseline, MIP, Average, Time-To-Peak, Peak-Enhancement, and the resulting curves)
result object | contains all necessary information for continuing an evaluation and printing the patient report
REV | Revision
RF | Radio Frequency
RFA | Radio Frequency Ablation
RFC | Request for Comments
RGB | Red-Green-Blue (color model based on additive color primaries)
RGMII | Reduced Gigabit Media-Independent Interface
RGSc | Respiratory Gating for CT Scanners (Varian)
RHem | Right Hemisphere
RHS | 'Right Hand Side' of the AQWP displaying on the right monitor the user interface of the embedded syngo.via software that is running on the IPS or in a virtual environment on the ICS.
RiLi | Ring Light
RIM | Risk Mitigation
RIMA | right internal mammary artery
RIS | Radiology Information System
RISK | combination of the probability of occurrence of harm and the severity of that harm
RJ | Recon Job
RLL | Right Lower Lobe
RM | Receiver Management
RM | Requirment Management
RMB | Right Mouse Button
RMC | remote control
RMeasy tool | Application to evaluate SAP RM analyzes
RMF | Risk Management Framework
RMII | Reduced Media Independent Interface
RML | Right Middle Lobe
RMmed tool | combination of SAP functions and a Microsoft Excel application to assist in obtaining, recording and managing restricted materials.
RMS | Root Mean Square
RM-tool | Requirements Management and Engineering Tool
RO | Radiation Oncology
RoHS | Restriction of the use of certain Hazardous Substances
ROI | Region of Interest
ROI | Region of Interest which can be circular, elliptical or irregular
ROT | Rotation
ROT mode | Special scan mode with one rotation at one position
Router | SiDaNet component with the features of a device but the additional capability to forward data traffic to other networks, for example Ethernet.
RöV | Röntgenverordnung (Germany)
RP | Requested Procedure (DICOM)
RPM | Revolutions per minute
RPP | Readings per Projection
rpp1 | no flying focal spot with quater detector offset (QDO)
rpp1x2 | z flying focal spot with quater detector offset (QDO)
rpp2 | flying focal spot with two focal positions (can be rpp1x2 or rpp2x1)
rpp2x1 | standard j flying focal spot with 1/8 detector offset (EDO)
rpp4 | combined z and j flying focal spot with 1/8 detector offset
RPS | Rotation Position Sensor
RQ | Requirement
RR | RR-Interval = time between two consecutive R-peaks (ECG)
R-R interval | Time duration between two consecutive R waves of the ECG
RRT | Rapid Results Technology
RS | Requirement Specification
RSC | Regional Support Center
RSC | remote scan control
RSCI | remote scan control interface
rsh | remote shell
RSS | Radio Standards Specifications
RT | Real Time
RT | Radiation Therapy
RT | Radiotherapy
RTC | Rich Thin Client
RTC | Real-Time Clock
RTD | Real Time Display
RTP | Radiation Therapy Planning
RTP | Radiotherapy Treatment Planning
RTR | Real Time Reconstruction
RTSS | Radio therapy structure set
RTTE | Radio & Telecom Terminal Equipment – 99/05/EG
RUH | Remote Update Handling
RUL | Right Upper Lobe
RUSP | Ready to Use Software Product
RV | Right Ventricle
RVA | Right Ventricular Evaluation
RWF | Reporting Workflow (IHE profile)
RWVM | Real World Value Mapping
RX | Receiver, part or signal of a receiving unit
RxPDO | CANopen receive PDO
S | Scaling factor
S&L | Standard and Law
S/PDIF | Sony/Philips Digital Interface Format
S/PDIF | Sony/Philips Digital Interface
SA | System Availability = 1 - ((downtime 1 + downtime 2) / operating hours)
SA | Short axis
SAB | Sales Advisory Board
SAC | Standardization Administration of China
SAD | Source to Axis Distance
SAF | Switch Arbiter Function, analyzes the outputs of the Switch Output Interfaces SOI of all other Switch Input Functions and moves the information in case of a match to the selected transmission FIFO for the timeslot.
SAF | Firmware task for SAFety
Safety Function | a function (of the system) which is potentially risky and therefore is relevant to safety, e.g. “motion”
Safety Mechanism | concrete measure implementing a Protective Function, e.g. “motion supervision”
SAFIRE | Sinogram AFfirmed Iterative REconstruction
SAI | Signal Application Interface, allows the manipulation of a subset of signals through the application via software.
SAM | Signal Access Mask, controls which bits of the signal vector could be accessed via hardware and by the software, running inside the application.
SAM | Semi Automatic Measurement
SAP-EDM | SAP - Engineering Data Management
SAR | Slipring Adapter Rotating
SAR | Segmentation And Reassembly.
SAR | Segmentation And Reassembly, describes the mechanism to transport long data fields with cells, whereby each cell transports only a small number of bytes.
SAR | Specific absorption rate
SAS | Slipring Adapter Stationary
sas | Segmentation And Reassembly.
SAS | Signal Access Sublayer, delivers the interfaces to access signals via software or hardware. Separates the signals to the hardware and application (softwarte) interfaces.
SAS | Serial Attached SCSI
SASCR | Signal Access Sublayers Register, part of the sublayers management interface.
SASIMCR | Signal Access Sublayer Interface Multiplexer Control Register, 4 register A to H in total, part of the sublayers management interface.
SASSR | Signal Access Sublayers Register, part of the sublayers management interface.
SATA / S-ATA | Serial ATA (computer bus technology for connecting hard disks and other devices)
Save Log | Save Log is a mechanism which canbe initiated by the user in order to create a ZIP file containing all relevant log files and traces of the system. In case of an error, the SaveLog can be retrieved from customer site by RDIAG. It is a valuable measure for post mortem error analysis.
SBV | Sensor Bias Voltage
SC | Secondary Capture - Modality independent DICOM image format. SC images contain sufficient information to identify the patient, date and institution, they do not contain any scaling, or other modality information, so are effective "just pictures".
Scan | The data acquisition / measurement process is here called a scan. The set of data readings acquired during a scan belong together.
Scan | For CT a scan is the technical procedure that is performed by the scanner between "x-ray on" and "x-ray off". A spiral range consists of a single scan whereas a sequence range typically consists of multiple scans. The scan itself has no direct representation as a domain entity.
Scan Begin Event | The transition from dummy readings to data readings.
Scan cycle | A scan cycle in ZigZag scan (scan with continuous table movement between two selectable table positions in a fixed cycle time) is from first to second and back to first table position.
Scan cycle time | The total time from scan initiation to next scan initiation. (see "standard scan cycle time")
Scan End Event | The transition back to dummy readings.
Scan FoV | Isocentric measurement area where CT projection data of 360 degree is available
Scan Protocol | A region and organ specific protocol that e.g. consists of a topogram and several sequence or spiral ranges
Scan Protocol | represents (a part of) a patient examination or a service examination. It consists of protocol steps (measurement, reconstruction, user instruction, quality step); a protocol step ‘measurement’ is a range or an autorange; a reconstruction step may consist of multiple reconstruction jobs
Scan Protocol Assistant | Tool for generalized scan protocol handling in excel like layout
Scan Protocol Designer | Tool for generalized scan protocol handling in Excel like layout.
Scan to scan time | Applicable to Sequence CT only: The total time from scan initiation to next scan initiation including table feed.
ScanUI | User interface of scan and reconstruction functions
SCAP | Security Content Automation Protocol
SCC | Scan Control
Scheduled Workflow | The Scheduled Workflow Integration Profile establishes a seamless flow of information that supports efficient patient care workflow in a typical imaging encounter. It specifies transactions that maintain the consistency of patient information from registration through ordering, scheduling, imaging acquisition, storage and viewing. This consistency is also the foundation for subsequent workflow steps, such as reporting. Systems involved in this profile are: - Enterprise-wide information systems that manage patient registration and services ordering (i.e.,admit-discharge-transfer [ADT ]/registration system and hospital information system [HIS ]) - Radiology departmental information systems that manage department scheduling (i.e. radiology information system [RIS ])and image management/archiving (i.e. picture archiving and communication system [PACS ]) - Acquisition modalities
SCI | System Control Interface
SCM | Supply Chain Management
SCONA | System Controller with SiDaNet (SPECT)
SCR | Servo Controller Rotation
SCS | Service Configuration System
SCSI | Small Computer System Interface
SCT | Servo Controller Tilt
SD | Standard Deviation
SD | Secure Digital Memory card (SD card)
SD | StanDard
SDB | Switching Database, memory inside the DSF which contains the correlation between incoming DCI of a cell and DCI plus port and timeslot information for the outgoing cell.
SDB | Service Diagnostic Board
SDC | System data collection format
SDD | Software Design Description
SDD | Simple Dose Detector
SDDO | Safe Digital Diagnostic Output
SDGUI | Software Distribution Graphical User Interface
SDHC | SD High Capacity
SDI | Safe Digital Input
SDN | Signal and Data Network
SDN | Signal Data Network (short abbreviation)
SDO | CANopen Service Data Object
SDO | Safe Digital Output
SDRAM | Synchronous Dynamic RAM
SE | Single Energy
SE | Systems Engineering
SE UI | SE Utilization Intelligence: a mechanism to track customer workflows and usage of features or system parameters on basis of event log entries from installed base
SEC | Supervisions, Environment and Cooling
SEDL | SeRex Driver / Library
SELV | Safety Extra Low Voltage
Seq | Sequence
Sequence | a combination of a number of rot modes with given timing and/or given table feed
Sequence CT | Scanning technique with single rotation and optional table feed between the scans.
SERDES | Serializer / Deserializer
SeRex | Segmented Slipring Receiver with PCIexpress Interface
SeRex | SEgmented Slipring Receiver with PCIexpress Interface (P58 Receiver)
Series | Series (= DICOM series) connects an amount of images within an examination with comparable parameters (e.g. with or without contrast application).
Service Pack | Service packs are software distribution kits that may be used for short-term bug fixes solutions or support of new hardware components in case of discontinuation.
service tasks | test or link to the CS service document, the service task is triggered out of the extended message text
Service Update | These are Bug-Fixes or performance improvements for a specific customer better known by CT as CPx and may not only be for one specific customer.
session | time between opening and closing of CaScoring
SEVERITY | measure of the possible consequences of a hazard
SFC | Sub Frame Counter, 5 bit counter to extend the 3 bit FPI-counter to a total 8 bit counter. He is used to manage the transmission of connections which share one FPI-value inside the DAS.
SFC | System Feature Concept
SFD | Start of Frame Delimiter
SFDA | State Food and Drug Administration
SFDE | Split Filter Dual Energy
SFF | Small Form Factor
SFTP | Secure File Transfer Protocol
SFW | Protection Module Specific FW
SG | Sliding Gantry
SGCB | Stationary Gantry Connection Box
SH | Shrink
SH | Siemens Healthineers
Shall | This requirement must be fullfilled.
SHI | Signal Hardware Interface, allows the manipulation of a subset of signals via hardware.
SH-No | Teamcenter-ID of the database object
Should | This Requirement should be fullfilled but not must.
Shu | Shuttle
SI | System Integration
SiCu Box | SiDaNet Central Unit Box
SiDaNet | Signal and Data Network
SiDaNet | Signal Data Network (long abbreviation)
SiDaSim | SiDaNet Simulator: Software and PCI-Card to simulate a SiDaNet Node.
SIF | Switch Input Function. Entry function for the Data Switching Sublayer, contains the DSF an decides to which destination ports and in which timeslots the cell must be sent.
Signal | Real time capable data type with the minimum information length of 1 bit. The signals were transported inside each cell independent of the connection, to which the cell belongs.
Signal Group | Bundling of (binary) signal bits to groups, inside which the bits could be multiplexed in a SiDaNet Switch.
Signal profile | Describes the total number of (binary) signal bits, their bundling to bit groups and the multiplexing to the signal bits transmitted inside each cell.
Signal vector | Total number signal bits which originate in the multiplexing of the signals like described in the signal profile.
SII | Switch Input Interface, input to the DSF, contains the 4 bytes data, the EED, the DCI and number of the receive port..
SII | SiDaNet Image data Interface
Simple Image and Numeric Report | The Simple Image and Numeric Report Integration Profile facilitates the growing use of digital dictation, voice recognition, and specialized reporting packages, by separating the functions of reporting into discrete actors for creation, management, storage and viewing. Separating these functions while defining transactions to exchange the reports between them enables a vendor to include one or more of these functions in an actual system. The reports exchanged have a simple structure: a title; an observation context; and one or more sections each with a heading, text, image references, and, optionally, coded measurements. Some elements can also be coded to facilitate computer searches. Such reports can be input to the formal radiology report, thus avoiding reentry of information. This integration profile involves both the department imaging and reporting systems and the hospital-wide information systems such as: - Review or diagnostics image softcopy display stations (stand-alone or integrated with a HIS, RIS, PACS or Modality) - Reporting stations (stand-alone or integrated with a HIS, RIS, PACS or Modality) - Report management systems (standalone or integrated with a HIS, RIS, PACS or Modality) - Report repositories (stand-alone or integrated with a HIS, RIS or PACS)
singapore | (former syngo) basic software platform for medical systems provided by Siemens Medical Department
single tomogram system | single-slice CT
SIP | Signal Input Part
SIS | Safety Integrity Supervisor
SiSo | Single Source
SIT | System Integration Test
Site Plan | The Site Plan contains all configured test cases which shall be executed in one test run. In the Site Plan the configured test cases are assigned to the sites where they shall be tested.
Site Planning | Assignment of test cases to real test systems including leveling of test system load.
SKU | stock keeping unit
SLC | System Load Computer
Slice | Image plane of a scan, defined by slice thickness and position.
Slice Mode | Image reconstruction algorithm without cone beam correction
Slice Mode | Image reconstruction algorithm without cone correction
Slice Number | Each slice block within a reading is identified with a 10 Bit number. The Slice number gives the physical number of the slice, which corresponds to it’s geometrical position.
Slice Start | A control in the received data stream indicating, that the next block received will contain reading data of the next slice.
slot | storage space for a data volume, e. g. the "CTA" slot for the CTA data. The term "slot" is used internally in this document
slot dialog | in the UI this dialog is called "data selection dialog". Here the user can select which data to display fused
SLR | Slipring
SLTT | single layer table top
SMBIOS | System Management BIOS (Basic Input Output System)
SMC | System Management Center
SMC | Switching & Monitor Control
SMD | surface-mounted device
SMI | SiDaNet Module Interface
SMI V2 | SiDaNet Module Interface defined for NPU
SMM | Signal Multiplex Matrix, operates the generation of the output signal bits as a function of all input signal bits.
SMP | SiDaNet Management Protocol: internal protocol for configuration, diagnosis and management purpose of SiDaNet
SMPTE | Society of Motion Picture & Television Engineers
SMS | Signal Multiplexing Sublayer, multiplexes the received signal bits inside each signal group to the transmitted bits.
SMSCR | Signal Multiplexing Sublayer Control Register, part of the sublayers management interface.
SMSDCR | Signal Multiplexing Sublayer MSI to SVM Delay Control Register, part of the sublayers management interface.
SMSPSR | Signal Multiplexing Sublayer Port Selection Register, part of the sublayers management interface.
SMSSR | Signal Multiplex Sublayer Status Register, part of the sublayers management interface.
SMSTCR | Signal Multiplexing Sublayer Taxi Postion Control Register, part of the sublayers management interface.
SN | Siemens Norm
SNAC | Siemens Nuclear Acquisition Controller (SPECT)
SNI | SiDaNet Network Interface, software communication interface located at switches for SiDaNet management
SNM | SiDaNet Network Manger
SNMP | Simple Network Management Protocol
Snow Med | Coding System for medical applications
SNR | signal to noise ratio
SNr. | Sachnummer = Siemens part no.
SOF | Switch Output Function, generates the output data for a port from the outputs of the SIF of all other ports.
SOI | Switch Output Interface, output from the DSF, contains the 4 bytes data, the EED, the new DCI and the Data Distribution Tag DDT.
SOM | SOMARIS
SOM | System Owner Manual
Somaris_Target | Somaris/7 system
SOMATOM LifeNet | Branding of Siemens CT extranet access
SOP | Service-Object-Pair, a DICOM Object Class
SOP | Signal Output Part
SP | Slice Position
SP_F | SPeaker at the Front side
SP_R | SPeaker at the Rear side
SP02 | oxygen saturation
SPC | Service Parts Catalog
SPC | Scan Protocol Control
SPCI | Standard Protocol of Camera Interface
SPDIF | Sony/Philips Digital Interface Format
SPDS | SPECT Power Distribution System
SPE | (Supply of insulated) signals for patient environment
Spec. | Specification
SPECT | Single Photon Emission Computed Tomography
SPI | Serial Peripheral Interface
Spi | Spiral
SPIRAL CT | Scanning technique with continuous rotation, continuous data acquisition and continuous table travel.
SPL | Sound Pressure Level: Is a logarithmic measure of the effective pressure of a sound relative to a reference value.
SPM | Scan Protocol Manager
SPO2 | Oxygen Saturation
SPP | Standard Packet Profile , standard encoding of the packet control structure PCS.
SPP | Spectral Post Processing
SPR | Stopping Power Ratio
SPrD | Scan Protocol Designer
SPrD | Scan Protocol Designer, which is a tool to define and modify protocols.
SPTNL | Spleen Peak Time Navigation Line
SQL | Structured Query Language
SQMP | Software Quality Management Plan
SQRt | Square Root
SR | Structured Report
sr | Steradian
SR | DICOM Structured Report
SRCV | Sample Rate Converter. SW Module for sample rate conversion.
SRF | Stop Report Function
SRI | ScanRecon Interface Server
SRoS | Stopreport over SiDaNet: Safety circuit mechanism provided by SiDaNet
SRP | Stop Report Protocol
SRPO | Siemens Reference Phantom - OSTEO CT
SRS | Siemens Remote Service
SRS | Slip Ring System
SRS | Smart Remote Services
SSCR | Synchronization Sublayer Control Register,. part of the sublayers management interface
SSCT | Single Source Computed Tomography
SSD | Surface Shaded Display
SSD | Solid State Disk
SSD | Source to Surface Distance
SSDE | Single Source Dual Energy
SSDE | Size Specific Dose Estimate (SSDE) according IEC 62985 (2019) is a dose magnitude based on CTDIvol and corrected for the actual patient size determined from the patient's CT-data
SSFDC | Synchronization Sublayer Frame Decrement Counter, part of the sublayers management interface
SSFDCSR | Synchronization Sublayer Frame Decrement Counter Status Register, part of the sublayers management interface
SSFIC | Synchronization Sublayer Frame Increment Counter, part of the sublayers management interface
SSFICSR | Synchronization Sublayer Frame Increment Counter Status Register, part of the sublayers management interface.
SSH | Secure SHell
SSI | Synchronous Serial Interface
SSID | Service Set Identifier
SSL | Secure Sockets Layer
SSO | Single Sign On
SSS | Switch Synchronization Sublayer, is responsible for the synchronization of the transmit data rate to the receive data rate inside a switch. It could work in master or repeating mode.
SSSDR | Synchronization Sublayer SyncDelay Register, part of the sublayers management interface
SSSR | Synchronization Sublayer Status Register, part of the sublayers management interface
SSW | Service Software
standard scan cycle time | time from one 360° sequence scan to the next one with type specific nominal rotation speed, including max. 10mm table feed and synchronous reconstruction of one image per rotation.
Standard Test Configuration | This type of Test Configuration is defined with Test Configuration variables and their values. A Standard Test Configuration requires at least a Product Definition. Each Standard Test Configuration counts as a single Configured Test Case.
STAT | Static
static | same for PET and SPECT as single-phase for CT: a single volume
Stationary Controller | hardware that hosts the eSW
STD | Scan time Display
STD_F | Scan Time Display at the Front side
STD_R | Scan Time Display at the Rear side
STEP | STEP-File is the most widely used data exchange form of STEP. STEP can represent 3D objects in Computer-aided design (CAD) and related information. The format of a STEP-File is defined in ISO 10303-21.
STIG | Security Technical Implementation Guidelines
STM | SiDaNet Translation Module (MI)
STO | Safe Torque Off
StopReportCircuit | Current circuit, which connects the STOP buttons for the operator
STS | Short Time Storage
STS | Short Term Storage
Study | Study (= DICOM Study) defines a set of series within an examination.
Study | A Study is a collection of one or more series of medical images, presentation states, and/or SR documents that are logically related for the purpose of diagnosing a patient (see Digital Imaging and Communications in Medicine (DICOM), NEMA).
SUI | SiDaNet User Interface, software communication interface located at devices for SiDaNet management
SUPPORTED EXPORT FILE-TYPES | csv; xml; wrk
SUPPORTED PUBLISH FILE-TYPES | PDF/A; Word, Excel
suspect | Suspect links are an indicator that other objects might be affected by the changed requirement.
suspend | to intermittently stop or hold an action, i.e. scanning, which can be continued again without delay.
SUV | Standardized Uptake Value
SV | Stroke Volume
SVA | Site Visit Avoidance = number of notifications without site visit / number of notifications
SVG | saphenous vein grafts
SVG | Scalable Vector Graphic
SVGA | Super Video Graphic Array
SVI | Stroke Volume Index
SVM | Signal Vector Merger, merges the 40 bit of the signal vector, transported with each cell, to the complete signal vector like specified in the signal profile.
SVS | Signal Vector Splitter, reverse function to the SVM. Divides the whole signal vector to the 40 bit long segments, which are transported inside a cell.
SW | Software
SW Update | Update packages introduce changes to the existing software better known to BU CR as Service Packs and may not be intended for only one customer but all. These changes may comprise performance optimisations, bug fixes, or new safety or security features etc. An Update can only be applied to already existing and installed software.
SW Update | Update packages introduce changes to the existing software better known to BL CT as Service Packs and may not be intended for only one customer but all. These changes may comprise performance optimizations, bug fixes, or new safety or security features etc. An Update can only be applied to already existing and installed software.
SW Upgrade | Upgrade packages enhance the existing software with new features.
SWA | Software Applications
Switch | SiDaNet component, which distributes data traffic from each port to the other ports. Each switch contains two ore more network ports.
SXGA | Super Extended Graphics Array
SY | Business Line Syngo
SyncMaster | Switch who delivers the clock reference for the whole SiDaNet
SyncRepeater | Switch, who is synchronized by a SyncMaster and delivers the synchronization for the associated devices.
SyncSlave | Device, which is synchronized by a SyncMaster or a SyncRepeater. A SyncSlave uses the synchronization information inside his sublayers but did not synchronize other components.
syngo | Software platform for medical devices syngo® is a registered trademark of Siemens AG
syngo d.via | departmental via is a multi-user client/server based syngo.via solution provided by Siemens Healthcare SY
syngo e.via | embedded via is a syngo.via in combination with an ICS
Syngo Imaging Workplace | Workstation specifically designed to work in a Siemens PACS environment
syngo via | Workplace provided by Siemens Healthcare SY
syngo.via | syngo.via is a multi-modality image processing system, which is developed under the lead of the Siemens Healthineers SY department and which is capable of multi-user client/server operation.
SysMon | System Monitor
SysOff | System Power State Off
SysOn | System Power State On
System Availability | Calculated value: 24 hours/365 days subtracted by the downtime caused by repair, maintenace and update (in percent) The time, during which the customer can use the system.
System Management Service | is the name of a software and service package of the Customer Service Division
t | Time
T_MOT | Tilt MOTor
T_SENS | Tilt SENSor
TA | Stands for Test Automation
TA DB | TA DB stands for Test Automation Database and represents the whole current test automation environment for Somaris/7 and Somaris/8.
Table Top | part of the patient table which allows move the patient automatically to any allowed position within the gantry
TAC | Time Attenuation Curve - The curve represents the time dependent HU values within a ROI of a CT dynamic perfusion scan
tag | A tag is a mnemonic for a requirement type.
TAS | Traffic Analysis Sublayer, integrated network analyzer who allows to record the data traffic and to stop the record after the detection (trigger) of specific, configurable, events.
Taskflow - Taskflow Template | The Taskflow is a subset of the departmental workflow supported by syngo.via. The Taskflow typically covers the indication specific part of a workflow. A Taskflow is defined by a taskflow template, in which the orchestration of clinical tasks is described.
TaskSwitchTime | Time span before radiation starts when the active range task will be selected automatically for image display
TAVI | transcatheter aortic valve implantation
TB | Terrabyte (10^12 Bytes)
TBD | To be defined
TBDE | Twin Beam Dual Energy scan
tbe | Difference between acquisition date and time of the last image used for generation of the baseline image and date and time of the injection start. (be: baseline end)
tbl | Acquisition date and time of the last image used for generation of the baseline image. (bl: baseline last)
tbs | Difference between the acquisition date and time of the baseline image and date and time of the injection start. (bs: baseline start)
TBW | Total Byte Written
TCB | Table Control Board
TCM | Tube Control Monitor
TCO | total cost of ownership
TCO | Tube Collimator
TCP | Transmission Control Protocol
TCP/IP | Transmission Control Protocol / Internet Protocol
TCS | Tube Control System
td | Delay between start of contrast fluid injection and start of scan. (d: delay)
TE | Test Engineering
teamplay | teamplay is a product of Siemens Healthcare SY which provides the customer the possibility to get various information about all the radiology modalities located in different departments of his institution.
Teleradiology | Transfer of digital information such as images, reports to a remote station via telephone lines or wireless.
TEMAX | Test Engineering, Management and Execution system
test case | A test case in software engineering is a set of conditions or variables under which a tester will determine whether an application or software system is working correctly or not.
Test Configuration Variable | This is an element provided by Microsoft Test Manager. This is used to provide Variables for Microsoft Test Manager Test Configurations. Conifiguration Variables can be used to define Test configurations.
Test Configuration Variable Value | Test Configuration Variable Values is a list of Values for a Test Configuration Variable.
test plan | The principle goal of test planning is to establish the list of tasks that, if performed, will identify all of the requirements that have not been met in the software. The main work product is the test plan.
test record set | All Test Records for a single requirement of a specific product combined to a document including a summary page.
test run | All test executions for a specific version of software under test.
test suite | A test suite is a collection of test cases that are intended to be used to test a software program to show that it has some specified set of behaviours. A test suite often contains detailed instructions or goals for each collection of test cases and information on the system configuration to be used during testing.
TF | Taskflow
TFS | Team Foundation Server
TFS DM Aligned Solution | TFS Defect Management Aligned Solution containing the DWIT plus additional tools
TFS team project | A collection of work items, tests and further artefacts used by a team
TFS team project collection | A organizational unit for team projects
TFS Test Point | Generally speaking: A TFS Test Point is a unique combination of TFS test case, TFS test suite, TFS test plan and TFS test configuration. Microsoft Description: Represents the execution state of a TFS test case in a TFS suite against a configuration as part of a TFS test plan.
TFT | Thin Film Transistor
TFT LCD | Thin Film Transistor Liquid Crystal Display
TG66 | Task Group 66
THD | Total Harmonic Distortion
TIA | Transimpedance Amplifier
TIA/EIA-644A | Is a telecommunications standard from the Telecommunications Industry Association (TIA), an offshoot of the Electronic Industries Alliance (EIA).
TiB | Terrabyte (2^40 Bytes)
Time Until Images | The total period of time, which the customer needs to wait after pressing the 'Recon' button until all the reconstructed images are available for assessment on the ICS
Time-to-Peak | Time to Peak in 1/10 s (seconds) at a pixel location (xi, yi) is defined as ten times the difference in date and time stamp of an input image and the baseline image such that the enhancement of the pixel value at a location (xi, yi) reaches its maximum in this input image over all the input images.
Time-to-Peak image | A postprocessed image of the same size as the input images where the value at each pixel (xi, yi) is a function of the Time-to-Peak value (within the input images) at that location.
TLC | Tube Load Computer
TM | Test Management
TMP | Testing at Manufacturers Premises (IECEE)
Token | Object on which authorization checks are made
TOMO | Tomogram
Tomogram | Cross sectional image which is generated by a rotational scan.
tool | the purchased requirements management tool CaliberRM with the configurations, customizations, the corresponding proceedings, file systems, and so on
Tool | (In a Clinical Task) Functionality used to perform a step. Tools may be used as common tools or as step-specific tools. Common tools are always presented to the user independent of the performed step whereas step-specific tools will be presented to the user only as long as the step is performed.
Top | Topogram
Top Support | part of the patient table which allows stepping between the patient table and the gantry, used in ICT mode
TOPO | Topogram
Topogram | Scan with fix tube position and continuous table travel that results in an overview image.
Topogram | A scan mode with a constant tube position and continuous table feed.
TP | Time Point (Single volume of a dynamic CT scan)
TPC | Transmit Power Control
TPM | Taxi Position Multiplexer, first multiplexer hierarchy inside the SMM.
TPM | Trusted Platform Module
TPN | Time Point Navigation
TPS | Treatment Planning System
TR | Transformer
TR | Technical Report, a Report published by the ISO
TRA | Transformer
TRA | Threat and Risk Analysis
TRA | Threat and Risk Assessment
TRACK | follow / tailer ELEMENTTYPEs via RELATIONTYPEs
Transformation | image volumes are transformed to new image volumes, (e.g. filtering, model calculation, ..)
TreeListView | The TreeListView is a GUI control that is able to present hierarchical data with details in columns.
TRRS | A Phone Connector type
TS | top support
TS | Technical Specification, a Specification published by the ISO
TSDE | Two Spiral Dual Energy
TSG | Trouble Shooting Guide
TSP | Test Specification
TTI | Table & Tilt Interface
TTL | Transistor-Transistor-Logic
TTT | TAC Table Tool
Tunnel length | part of the gantry which has minimal aperture, important for PET
Turbo Recon | Option to speed up the reconstruction process.
TX | Test Execution
TX | Transmitter, part or signal of a transmitting unit
TxPDO | CANopen transmit PDO
UART | Universal Asynchronous Receiver Transmitter
UDC | Voltage of Intermediate Circuit
UDI | Unique Device Identification
UDP | User Datagram Protocol (standard network protocol)
UEFI | Unified Extensible Firmware Interface
UEV | Unexpected Value
UFC | Ultra Fast Ceramic
UFD | Ultra Fast Detector (Counting)
UHR | Ultra High Resolution
UHRplus | Ultra High Resolution in phi and z direction, achieved with special hardware comb
UI | User Interface
UI | Update Instruction
UI | User Interface (means all accessible UIs unless other described!)
UI | User Interaction
UID | Unique Identifier
UL | Underwriters Laboratories Incorporated
UL | Upper left
UM | Upper middle
UM | Utilization Management
UMAR | Universal Master Rotating
UMAS | Universal Master Stationary
underlay | the opposite of overlay, the part of the image underneath the overlay
UPD | Updater
Update | system modification with improvement of performance in terms of higher reliability, productivity, user friendlyness, error corrections, without new functionality
Upgrade | An upgrade offers increased functionality to the existing system, i.e. new applications or new features for applications.
UPS | Uninterruptible Power Supply
UPS | Unified Procedure Step
Uptime | Availability of the system; it is calculated as Uptime [%] = 100 - downtime for partial and total failure/operating time
UR | Upper right
UR | Use Requirement
US | Use Scenario
USB | Universal Serial Bus
USC | Uptime Service Center
user | the person who use the system in its intended use (patient examination, reconstruction, data handling, etc.; e.g. technologist, radiologist)
User | a CT system has different unsers: clinical user (is the standard user), clinical admin, radiologist, IT Administrator, CT (as manufacturer), Service Technician
user profile | User profiles encompass the user-specific storage and reutilization of all settings accessible through the system.
USERDOC | Environment variable containing path of the e-manual content, e.g. U:\UserDoc -- use as follows: %USERDOC%
USI | User Interaction
USP | Unique Selling Point
USV | Unterbrechungsfreie Stromversorgung
UTC | Coordinated Universal Time Coordinated Universal Time is the primary time standard by which the world regulates clocks and time. Time zones around the world are expressed as positive or negative offsets from UTC.
UTT | Universal Test Tool
UVP | Under Voltage Protection
UWL | Unified Worklist
UX | All upper segments
UX | Design & User Experience Group
UX Guide | Windows User Experience Interaction Guidelines
V1 | first released version of ColonCT
V2 | second released version of ColonCT
VA | Veteran Administration
VAI | Variable Application Interface, allows the application (software) access to variables. Variables are identical to the data bytes which are transported inside one cell.
VALUE | stored information of an ATTRIBUTETYPE belonging to an ELEMENT
VAR | Volume Artefact Reduction
Variable | Real time capable data type with a constant information length of 4 bytes. The interpretation of a variable is done by the connection, inside which it is transferred. The QoS, which is required for a connection, could be adjusted by the data traffic management.
Varian RGSc | Varian Respiratory Gating for Scanners
Varian RPM | Varian Real-time Position Management
Vario1 | Table with vertical movement for Project P81B
VCR | Video Controlled Recording
VCT | Volume CT
VD | Virtual Dissection
VDE | Verband deutscher Elektrotechniker
VDI | The Association of German Engineers
Vector | Table without vertical movement for Project P81A
VESA | Video Electronics Standards Association
VGA | Video Graphic Array
VGT | Virtual Gantry Tilt
VHI | Variable Hardware Interface to access the variables via hardware.
VIA | string used to mark a series describtion used for syngo.via
VIBE | Volumetric Interpolated Breathhold Examination
VLA | Lertical long axis
VNC | Virtual Non Contrast
VOI | Volume of Interest
Volume Mode | Image reconstruction algorithm with cone correction
Volume Recon | Reconstruction processing due to volume resolution and cone beam correction.
VoV | Volume of View
Voxel | Volume element of a image, extension of pixel for three dimensional considerations.
VPCT | Volume Perfusion Computed Tomography
VPCT | Volume Perfusion CT
VPI | Visual Patient Instruction
VPN | Virtual Private Network
VRF | Variable Transmit Function, controls the handling of a received cell with variable data inside the DAS.
VRI | Variable Receive Function, controls the receiving of variables inside the DAS.
VRT | Volume Rendering Technique
VSIM | Virtual Simulation
VSP | Variable System Parameter
VSTS | Visual Studio Team System
VTF | Variable Transmit Function, controls the generation of a cell with variable data for transmission inside the DAS.
VTI | Variable Transmit Function, controls the transmission of variables inside the DAS.
VVS | Variable Velocity Spiral
VZ | Volume Zoom
WA | Worklfow Assistant (optional documentation)
WAA | Wireless Access Device as Accessory
WAD | Wireless Access Device
Wall Switch | an optional module, e.g. Sirius, will be added to the on-site installation, to disconnect the system from mains when switched off. Interfaces will be provided to implement a safe startup and shutdown, triggered by “On/Off Button” as well as from the on-site installation
WAN | Wide Area Network
warm start | start from hibernate state until system is ready for measurement with specified image quality.
WBT | Web Based Training
WCF | Windows Communication Foundation
WCS | Water Cooling System
WEEE | Waste of Electrical and Electronic Equipment
WHO | World Health Organization
WI | Workitem (entity of clustered OCTs which is used by a user at one time)
WiFi | Wireless Fidelity
WIG | Wireless Interface Gantry
WiGig | Wireless Gigabit
WII | Wireless Interface ICS
Win32 API | Application Interface for 32-bit Windows Applications
WIP | Work in Progress
Wizard | marketing name for IES (shall not longer be used as abbreviation)
WLAN | Wireless Local Area Network
WMI | Windows Management Instrumentation
WMSY | syngo
WMV | Windows Media Video
WOL | Wake on LAN
WORKFLOW | is a set of sequential activities and defined status which is reflected as user activities, integrated in the RM- tool UI
Workflow | Workflow is the term to describe a work sequence to be followed in syngo.via with clinical work, e.g. reading cases. A workflow consists of one to many tasks. Taskflow is a synonym to Workflow.
Workflow | represents the complete course of actions that are performed during an examination is the upmost level of the hierarchy workflow - patient - study - scan protocol
Workflow - Workflow Template | The workflow is a subset of the departmental workflow supported by syngo.via. The workflow typically covers the indication specific part of a workflow. A workflow is defined by a workflow template, in which the orchestration of clinical workflow steps is described.
Workflow Duration | The time elapsed from user interaction initiating the first procedure in the workflow (e.g. starting topogram acquisition), to the moment, when all reconstructed images are available for assessment on the ICS
Workflow Step | A workflow step represents a clinically meaningful and self-contained piece of work performed by one user. Usually multiple workflow steps are performed in a specific order to answer a clinical question. Workflow steps may be assigned to different roles.
Workplace | Genus for a CT console – includes AWP, CTWP and MMWP
WPA | Wi-Fi Protected Access
WPA2 | Wi-Fi Protected Access 2
WR | Wechselrichter
WS | worksheet
WS | Wall Switch
WSpM | Weighted Subpixel Map
WWW | World Wide Web
WXGA | Wide Extended Graphics Array
wysiwyg | what you see is what you get
x | placeholder for variable
XACB | eXtended Access Control Block
XC | Synchronization signal to trigger X-ray scan
XDC | X-ray Deflection Control
XGR | X-ray Generator Rotating
XGS | X-ray Generator Stationary
XML | Extensible Markup Language
XP | Windows XP -Operating System
XRS | X-ray System
XRS | X-ray system
XRT | X-ray Tube
XTA | X-ray Tube Assembly
XTC | X-ray Tube Cooling
XTM | X-ray Tube Memory
XXL | XXL patients instead of obese patients
YY | chinesischer Standard für den Industriesektor der Medizingeräte
YY | Medical Industry Standard (China)
YY/T | YY/tuījiàn, empfohlener chinesischer Standard für den Industriesektor der Medizingeräte
ZEC | Z-axis Exposure Control
Zig | Zigzagscan
ZigZag scan | Scan with continuous table movement between two selectable table positions in a fixed cycle time.
ZM | Zeichnungsmuster
ZMT | Application in SAP
Z-sharp | Ultra high resolution mode using the phi-z-com
zTest_do_not_use | test